## Slide 1: CPIXTEG3b: An SOI pixel sensor with in-pixel binary counter
Y. Lu1, R. Hashimoto2, S. Kishimoto2, R. Nishimura3, L. Song1,4, Y. Zhou1, Z. Wu1,4, Y. Arai5, Q. Ouyang1
1 State Key Laboratory of Particle Detection and Electronics, Institute of High Energy Physics, CAS
2 Institute of Materials Structure Science, KEK
3 School of High Energy Accelerator Science, SOKENDAI
4 University of Chinese Academy of Sciences
5 Institute of Particle and Nuclear Studies, KEK
International Conference on Technology and Instrumentation in Particle Physics,
May 22-26, 2017, Beijing, CHINA
Inter-University Research Institute Corporation
High Energy Accelerator Research Organization

## Slide 2: Outline
Motivation
  To solve digital pickup issue in SOI technology
  To pursue a high resolution low noise detector for material science
Design concepts
  Diode, Pixel circuit, Readout architecture
  Double-SOI oriented optimization
  Test structures
Single pixel test
  Analog signal inspection
  Dedicated pickup measurement
  Infrared laser test
Matrix test
  Threshold
  Noise
  Synchrotron beam test
Conclusion
TIPP2017, Y. LU

## Slide 3: Motivation #1: to solve the digital pickup issue
SOI pixel technology is particularly affected by the digital pickup issue.
Long history of investigation, but limited by the process available then.
  recognized and understood by CERN RD19 in 1990’s;
  Continued efforts on Nested-Wells in the beginning of 2010’s.
Shielding-well proposed by F. X. Pengg in his dissertation “Monolithic Silicon Pixel Detectors in SOI Technology”
0.2kΩ/□
9.4kΩ/□
Monolithic Active Pixel Matrix with Binary Counters ASIC with nested wells
F. Fahim, Pixel2012
BOX,
Buried OXide
G. DEPTUCH, SOIPIX 2010
~10kΩ/□
TIPP2017, Y. LU
3

## Slide 4: Motivation #2: high resolution low noise detector for the material science
New ferroelectric materials study requires measurement of the intensity of specific diffraction spots with
  30um square pixel;
  1k frame/s;
  14-bit counter each pixel;
Determination of structure change of cell membrane
  Grazing-incidence small-angle X-ray scattering (GISAXS);
  Minimum area 20~30mm2;
  2.1~4.5keV, thin entrance widow < 1um;
TIPP2017, Y. LU

## Slide 5: Sensor diode
  > 1kΩ-cm high resistive substrate
  P type
  300um thick
Charge collection electrode
  N type
  Consists of NS and BNW
NS
SOI Si
Buried
Oxide
(BOX)
Cut Top Si and BOX
High Dose
Substrate Implantation
BNW Implantation
BNW
Keep Top Si not affected
Low Dose
Pixel
Peripheral
N-in-P sensor
  Peripheral BNW as single GR
  P-stop ring needed (PS implant)
TIPP2017, Y. LU

## Slide 6: Pixel circuit
PREAMP
SHP
Counter
Charge sensitive preamplifier
  Cf = 4fF
  Constant current feedback (0.5 - 10nA adjustable)
Shaper with inverse polarity
  AC coupled
  Voltage gain = 5
Diode-biased inverter as the discriminator
  AC coupled
  Global Ithreshold tuned by a 4-bit local DAC
6-bit ripple counter
  Switchable to shift register mode for readout;
  Counting depth compromised to achieve a small pixel pitch of 50um;
6-bit shift register
  For control of each individual pixel;
  DAC setting, Pixel mask, Calibration enable.
SOI2 contacts
Signal processing chain in-pixel
TIPP2017, Y. LU

## Slide 7: Readout architecture
PREAMP
SHP
Counter
Counters and registers are accessed by a 6-bit data bus
  64*64 pixels
  Driven by a 20MHz clock
  Readout time ~ 200us/frame
SOI2 contacts
Data chain organized in column
Pixel(0,0)
Pixel(0,62)
Pixel(0,63)
…...
TIPP2017, Y. LU

## Slide 8: Double-SOI process
SOI2-related Capacitance and sheet resistance
Advantages:
  Back-gate under control;
  Flexible scheme of shielding;
  Compensation to the trapped charge in BOX (SiO2);
Disadvantages:
  Capacitance between BNW and SOI2 (C2 ~ 0.23fF/um2);
  High sheet resistance of SOI2 (R1 ~ 30kΩ/□).
~ 30kΩ/□
SOI2
TIPP2017, Y. LU

## Slide 9: Critical design choices to enhance shielding
The collection electrode kept clear of counter/discri., which leaded to
  A small collection electrode of 16 um to reduce diode capacitance
  A small pixel pitch at 50 um with a concerning on CCE
The shield grounding improved by
  As many SOI2 contacts (2CT) as possible
  Local bypass capacitor on SOI2 ground
P-stop ring isolating pixels
  SOI2 island formed by PS cutting
Pixel layout
TIPP2017, Y. LU
SOI2
Double-SOI cross-section view

## Slide 10: Test structures
  3 test pixels with different configuration:
  Pixel(0,0), disconnected from sensor, frond-end only
  Pixel(0,31), a complete chain of signal processing
  Pixel(0,63), an alternative clock to its counter
Chip layout
TIPP2017, Y. LU
Pixel(0,0)
Pixel(0,31)
Pixel(0,63)

## Slide 11: Analog signal inspection
Double-SOI chips worked
  Slower leading edge was caused by improper cabling;
In the contrary, Single-SOI chips showed self-sustained oscillation
  Managed to stop it by decreasing gain of shaper
  A clue of positive feedback loop
X 1/5
X 1/100
128 frames averaged
Self-sustained oscillation observed on single-SOI chips
Preamp Shaper
Sensor
TIPP2017, Y. LU
Waveform observed on oscilloscope (average mode)
Pixel(0,31)

## Slide 12: Dedicated measurement of pickup from counter
Counter driven by external clock as a source
5mV @ shaper output for Double-SOI chip
  74 e- referred to input charge
  Submerged in noise floor (ENC ~ 52e-)
95mV for Single-SOI chip
  ~3770e- referred to input charge
128 frame averaged
X 1/5
(peak to peak) | Double-SOI | Single-SOI
Preamp output | 3.7mV | 60mV
Shaper output | 5mV | 95mV
TIPP2017, Y. LU
Waveform observed on oscilloscope (average mode)
Pixel(0,63)

## Slide 13: Laser test
Infrared laser beam
  1064nm, simulating MIP tracks in silicon;
  Sub-ns pulses duration and maximum 1MHz;
Signal amplitude increased as a function of Vbias
  Proportional to ;
Counting rate decreased as the threshold lifting
  S-curve fitting resulted in 153e-;
  1676 e- generated by single laser pulse
TIPP2017, Y. LU
Pixel(0,31)
Preamp output waveform (us)
Amplitude vs Vbias
Counting rate vs input-referred threshold

## Slide 14: Threshold distribution
Test pulse amplitude scan
  S-curve fitting
  3966 pixels
Excellent threshold distribution
  44.6e- before tuning
  10.8e- after tuning
Threshold dispersion after tuning
Threshold dispersion before tuning
Typical S-curve fitting
44.6e-
10.8e-
TIPP2017, Y. LU

## Slide 15: Low noise level
Average noise 56.7e-
  3966 pixels
“Zero” noise hit
  Threshold = 830e- (3keV)
  10s ×360 frames (for long exposure time and small data set)
  Only a few events recorded from environmental radiation
Environmental background events in 1 hour
Col. 0
disabled
Pixels broke up for test
52.3e
TIPP2017, Y. LU
Noise distribution

## Slide 16: Environmental background events (10s ×360 frames)
TIPP2017, Y. LU

## Slide 17: Synchrotron beam test
KEK PF BL-14A
  Beam spot 0.8mm in diameter
  optional pin hole, 10um micro-beam
16 keV X-ray signal(shaper)
(16keV, threshold~1600e-)
65419 photons
16×19 pixels
Beam profile measured by CPIXTEG3b
TIPP2017, Y. LU

## Slide 18: Micro-beam 1D scan
Scan across 4 pixels from center to center
  Starting at the center of Pixel(30,37)
  Ending at the center of Pixel(30,41)
  Step = 2um
  NaI + MCA used as a reference detector
Uniform efficiency can be achieved
  Threshold = ½ Ephoton
16keV, threshold > ½ Ephoton
16keV, threshold ~ ½ Ephoton
8keV, threshold ~ ½ Ephoton
6keV, threshold ~ ½ Ephoton
Pixel(30,37)
Pixel(30,41)
Pixel(30,39)
sum
Individual
Pixel(30,38)
Pixel(30,40)
TIPP2017, Y. LU
1-dimentional scan across 4 pixels

## Slide 19: Micro-beam 2D scan
Across 3×3 pixel-square
  Step = 4um
Inevitable loss of efficiency at the corner
  Threshold = ½ Ephoton
  Charge shared by 4 pixels adjoined
2-dimentional scan of pixel corners
6keV
16keV
Pixel corner
TIPP2017, Y. LU

## Slide 20: Flat field response
Beam scattered by glassy carbon
  Sensor placed in 90⁰ with respect to the beam line
Full pixel array illuminated.
  “Relative uniformity” 3.8% and 2.0%
Mean:9382.5
RMS:183.0
2.0%
Mean:962.9
RMS:36.9
3.8%
6keV
16keV
Counting rate histogram
TIPP2017, Y. LU

## Slide 21: Conclusion and outlook
Digital pickup issue can be treated properly with Double-SOI process
Further improvements to meet the needs of material science
  Charge sharing decision logic;
  Compact layout of counter to accommodate more bits.
CPIXTEG3b demonstrated the potential of SOI technology
  Zero noise hit
  Excellent threshold distribution
  Depleted HR substrate up to 700um (or thinning down to 75um)
  High density connection between sensor and electronics
TIPP2017, Y. LU