## Slide 1: SiPM camera R&D for LHAASO-WFCTA
Shoushan Zhang for LHAASO collaboration
International collaboration: Geneva University, INFN
Institute of High Energy Physics, Beijing, China
TIPP 2017, May 22-26, 2017, Beijing

## Slide 2: Outline
Introduction
LHAASO-WFCTA SiPM camera R&D
  SiPM camera design introduction
  Performance of SiPM candidates
Conclusion and discussion

## Slide 3: Why choose SiPM?
Advantages: single photon response, high detection efficiency, high gain at low bias voltage，no aging due to light exposure and no sensitive to magnetic fields.
Because SiPM not aging due to strong light exposure, SiPM-based camera can be operated in moon conditions, thus achieve a longer exposure time than PMT-based camera. The exposure time of SiPM-based camera is about 30%, while the PMT is only 10%.
The first G-APD Cherenkov telescope (FACT): La Palma, Canary Islands, Spain，2011
CTA: High energy Cherenkov telescopes (Single mirror small size telescope), Schwarzschild-Couder telescope
A SiPM camera prototype of CTA
FAST

## Slide 4
WCDA:78,000 m2 3000 cells (25m2/cell)
WFCTA:
16 Cherenkov telescopes (1024 pixels/telescope)
LHAASO
KM2A:
5195 EDs
1171 MDs
Daochen, Sichuan (29o21’ 31” N, 100o08’15” E, 4410 m a.s.l., 600 g/cm)

## Slide 5: Wide Field of View Cherenkov Telescope (WFCTA)
SiPM camera
5m2 spherical mirror；
32×32 SiPMs array
Pixel size 0.5°；
FOV: 14°× 16°.
A prototype of WFCTA
instead of

## Slide 6: LHAASO-WFCTA vs. CTA
 | CTA-SSTs | LHAASO-WFCTA
Energy range | 5 TeV – 300 TeV | 30 TeV – 1 EeV (Cherenkov mode: 30TeV – 10 PeV, 10PeV-100PeV; Fluorescence mode: > 100PeV)
Diameter of mirror | ~ 4 m | ~ 2.3 m
Pixel size | ~ 0.25 | ~ 0.5
Number of pixels | 1296 | 1024
Dynamic range | 1 pe - 2000 pe | 10 pe – 32000 pe
Light pulse duration | 6 ns - 50 ns | 6 ns - 3 μs
Observation mode | Cherenkov light | Cherenkov and Fluorescence light
Fluorescence light spectrum
Air shower Cherenkov light spectrum

## Slide 7
 | SiPM | PMT
Dynamic range | Proportional to the total number of APD cells | Dependent on PMT type, HV distribution and readout method
Photon counting | Yes | No
Charge resolution | 12-13%@100pe，4%-5%@1000pe | 12%@100pe, 4%@1000pe
Pulse duration | 20 - 60 ns | ~6 ns
Dark count rate | ~80 kHz/mm² @ thrd=0.5 pe | <1kHz @ thrd=0.5 pe
Aging@strong light exposure | No aging | Aging
Magnetic fields | No sensitive | Sensitive
Gain | 1×106 @ several tens volts | 1×106 @ >1000 V
Gain vs. temperature | ~1.5%/℃ | ~0.2%/℃
LHAASO-WFCTA prototype (PMT array)
PMT aging over time

## Slide 8: Schematic diagram of SiPM camera assembly
AB1
AB2
DB
Cluster
SiPM camera box
SiPM camera：32×32 SiPMs
64 ×
A module has 4×4 SiPMs.
Optical filter

## Slide 9
Digital Board
Analog Board
Slow control board
Metal frame
SiPM board
Winston cone
Light concentrator (Winston cone): is to increase the effective area of the SiPM
SiPM board: 16 SiPMs, 16 temperature sensors (embedded in the SiPM chip ), and 16 pre-amplifiers
Slow control board:
  16 temperature and high voltage compensation loops;
  16 channels of high voltage power supply.
Analogue board: 16 channels of amplifier and shaping circuit
Digital board: 50 MHz FADC and FPGA
  Square SiPM：15 mm×15 mm

## Slide 10
Outlet: 15 mm×15 mm
Inlet: 24.4mm×24.4mm
Height: 25.3mm
  Square SiPM：15 mm×15mm
Pixel size: 25.4 mm × 25.4 mm
SiPM size: 15 mm × 15 mm
light concentrators
0°
8°

## Slide 11
4 × 4 SiPMs & Preamplifiers
Service
TCP/IP
Switch
Low voltage power supply
Slow Control Board
64 ×
DB
FPGA
50M 14 bit FADC
50M 14 bit FADC
BDB
Trigger & Data Buffer & Communication Board
Event trigger
Data
AB
Low
gain
High
gain
Synchronizing
TCP/IP
The SiPM gain (break down voltage) is sensitive to the temperature.
Slow Control Board: temperature & high voltage compensation loop.
16 ×
Temperature
High
voltage
The architecture
of SiPM camera
readout electronics
Single trigger

## Slide 12
R4
-
+
vcc
SiPM
C1
R1
R2
OPA 846
R3
Pre-amplifier
A prototype of SiPM board
A prototype of AB and DB board

## Slide 13
Sky background light spectrum
in the moon less night

## Slide 14
Chart: 
Chart: 
Filter window：AR + cutoff

## Slide 15: SiPM candidates: HAMAMATSU, SensL and FBK
 | FBK (25 μm) | SensL (25 μm) | Hamamatsu （25 μm）
PDE | 38%@400 nm | 31%@400 nm | 25%@400 nm
Fill factor | 72% | 62% | 47%
Dark count rate | ~80kHz/mm2 | ~50k/mm2 | ~45k/mm2
Optical cross talk | ~15% | ~5% | ~1%
Gain | ~2×106 | ~0.9×106 | ~0.7×106
Break down voltage | ~26.5V | ~24.5 | ~53V
Gain vs. temperature | ~1.5%/℃ | ~1.5%/℃ | ~1.5%/℃
Break down voltage vs. temperature | ~26 mv/℃ | ~24 mv/℃ | ~54 mv/℃

## Slide 16
Low light intensity response
HAMAMATSU ：12 mm×12mm
Low light intensity response
FBK：5mm×5mm
FBK
HAMAMATSU
S13361-3050AS-04
FBK SiPM testing board
HAMAMATSU SiPM testing board

## Slide 17
  HAMAMASTU: 12 mm ×12 mm
Ncell =230,400
The dynamic range of the SiPM is proportional to the total number of cells.
  Larger area SiPM has a bigger DCR.
  The small APD size: will produce small fill factor and then small PDE.
  SiPM with photosensitive area=15mm×15mm and APD size=25μm is selected.
FBK: 25mu
12×5 mm×5 mm
Ncell =480,000

## Slide 18
Light spot from the reflective mirror
Estimate the non-linearity from the non-uniform light distribution cause by the light spot and the light concentrators
Light distribution output from the cone
~1.5% @ 32000 pe

## Slide 19
PMT
FBK
Charge resolution:
  SiPM: ~5%@1000pe
  PMT: ~4%@1000pe
  HAMAMASTU
SensL

## Slide 20
The relative deviation is less than 3% with C1=1 μF.
The long light pulse response
C1=1 μF
  HAMAMASTU

## Slide 21
 | FBK | SensL | Hamamatsu | PMT （R7899）
PDE | 38%@400 nm | 31%@400 nm | 25%@400 nm | 20%@400 nm
Dark count rate | 18 MHz | 11.3MHz | 11.3MHz | <1kHz
Sky background noise | ~49 MHz | ~40 MHz | ~32 MHz | ~24MHz
Signal to noise ratio
SiPM: 15mm×15mm，25μm
1.13 (FBK), 1.06 (SensL), 0.93 (Hamamatsu)

## Slide 22: Conclusion and discussion
SiPM camera are designed and can meet the requirements of LHAASO-WFCTA
  SiPM-based camera can be operated on moon nights and achieve a longer duty cycle than PMT-based camera.
    -- Energy threshold of the telescope: ~30 TeV @ moonless night, ~300 TeV @ full moon night.
    -- SiPM based camera and PMT based camera almost have the same S/N at a moonless night.
  Time schedule
    -- First SiPM camera will be built before Jan. 2018;
    -- Six telescopes based on SiPM camera will be built before the end of 2018.
SiPM candidates: Hamamatsu, FBK and SensL SiPM can meet the requirements of LHAASO-WFCTA

## Slide 23: Thank you for your attention!

## Slide 26: I-V curve for breakdown voltage and quenching resistor
Linear fit
Breakdown voltage: 53.3 mv/℃
Gain: ~1.5%/℃

## Slide 27
20μm
25μm
20μm
25μm
Breakdown voltage vs. temperature: ~26 mv/℃
Gain vs. temperature：~1.5%/ ℃

## Slide 28
Moon trajectory
North
East
South
West

## Slide 29
20μm
25μm
20μm
25μm
Breakdown voltage vs. temperature: ~26 mv/℃
Gain vs. temperature：~1.5%/ ℃