## Slide 1: The SLAC Instrumentation and Control Platform
R. Claus, M. D’Ewart , J. Dusatko, R. Herbst, K. Kim ,
L. Ma , U. Legat, J. Olsen, B. Reese, L. Ruckman,
L Sapozhnikov, S. Smith, T. Straumann J. Vasquez
D. Van Winkle , M. Weaver, E. Williams, C. Xu, A. Young

## Slide 2: The ATCA Common Platform
Developed at SLAC for high performance distributed control systems
Based on the Advanced Telecommunications Computing Architecture (ATCA)
  Growing multi-billion dollar market
  Continuously improving technology
High density, backplane carries all global signals
ATCA Crate “shelf”
Standard hardware, SLAC application:
Available from 1 slot to 16 slot
Power supply with hot swap
Management
Networks
Cooling
Designed for high-reliability applications

## Slide 3: Common Platform
Carrier card: FPGA, memory, backplane connections
AMC cards: ADCs, DACs, high performance front end electronics
RTM: General purpose IO, extra networks, miscellaneous
1: Carrier Card
2: Crate
3: RTM
Each carrier supports 2 AMC application cards
SLAC hardware and software application on ATCA

## Slide 4: ATCA Backplane (As used in SLAC Common Platform)
Ethernet for data / control (10Gbit, 10/40 uplink)
Timing data stream (3.75Gbit)
Machine protection / uplink (1.25Gbit)
Slot 1
switch
Slot 2
payload
payload
payload
payload
payload
payload
payload
10G / 40G data
IPMI
MPS
Timing
Single cards can be used with rear transition module networking
All carrier cards can also serve as slot 2 nodes
Crates available from 1 to 16 cards
IPMI management network
For most applications, all networks are carried on the backplane, minimal external connections

## Slide 5: More on Backplane Networks
IPMI provides control to power cycle and hot-swap individual cards.
  Operates independently of FPGAs on carrier cars.
Ethernet network used for normal control and data.
  10Gbit to each card from switch
  10 or 40Gbit uplink to main network
Timing data stream provides synchronization
  Each carrier card decodes the data stream to generate internal or external triggers. Multiple data streams supported:
    Micro-Research “event receiver”
    SLAC LCLS2 “timing system”
    Straightforward to support other serial timing systems
Machine Protection / uplink
  Used at SLAC LCLS2 to provide an independent uplink path for interlock data.
  Available for other real-time uplink use if MPS is not required
  For typical applications there is one set of networks connected to the crate, but no networks to the individual cards

## Slide 6: Carrier Card
AMC ADC/DAC cards
Connector to backplane
FPGA
Connector to RTM
Power connector
DRAM
Power supplies for AMCs

## Slide 7: Carrier Card
Xilinx “Kintex Ultrascale” FPGA
  KCU040 or KCU060: 2760 DSP slices, 200Gbit bidirectional bandwidth for ADC / DACs
Power supplies for AMC cards
  +3V, +4V, +/- 6V, +/-15V, +12V (9A)
  No switching supplies required on AMCs for most applications
8GB DDR3 memory (in addition to 3MB on-chip fast memory)
Supports Common Platform networks and management
Carrier card is very complex but a single design is used for all applications

## Slide 8: AMC Cards
AMC = Advanced Mezzanine Card
ATCA double width, full height AMC card.
  Fully backwards compliant with AMC standards, with additional features
  Full height card provides space for RF shielding and daughter cards
Combined ADC / DAC and analog front end
  ADC/DAC requirements usually tied to a specific front end design
  Analog engineers can work independently of digital engineers
A variety of AMC cards have been developed
  ADC/DAC, Loopback test, Beam position monitor, Parallel Digital IO, High speed serial IO, Precision trigger generator. Low level RF system, high performance ADC/DAC
Some example AMC cards shown on the next slides

## Slide 9: General Purpose A-D / D-A card
4X 16 bit 370 Ms/s ADC
  TI ADC16DX370
  One channel for read back of timing system for precision synchronization (sub-sample)
2X 16 bit 370Ms/s DAC
  TI DAC37J82
ADC/DAC clock
  DAC driven VCXO for locked or unlocked operation
  Low phase noise clock oscillator
Fast trigger reference output
2X DIN, 2X DOUT
Fiber output for triggering
Status: In Production
Larry. Ruckman
Used for 300nm resolution cavity BPMs

## Slide 10: Stripline e-Beam Position Monitor
Andrew Young
RF shields removed for picture
2 micron resolution at 180pC
BPM signal processing
  300MHz, 30MHz BW filters
  Distributed amplification /attenuation for high linearity
  370Ms/s ADC
Calibration function
  Inject signals on BPM strips
  Measure induced signal in perpendicular strips
  Corrects for cable and electronics drifts
Status: In Production

## Slide 11: Low Level RF Controller
2 AMC cards on a single carrier
  (receiver shown).
10X RF input channels
1X RF output channel
2X fast DC ADC (370Ms/s)
300MHz to 3GHz range with different oscillator daughter cards
Clocks and LO phaselocked to timing reference system.
FPGA provides closed loop RF control.
Very low measurement noise
  4.3 femtoseconds RMS in 1MHz bandwidth
  1 femtosecond RMS in 1KHz bandwidth
Status: Operational
Dan Van Winkle

## Slide 12: High Speed A-D / D-A
4X 2.5Gs/s 14 bit ADC, 4X 2.5Gs/s 16 bit DAC
200Gbit/s raw ADC / DAC bandwidth, reduced in ADCs / DACS to 100Gbit/s to the FPGA
Designed for cryogenic sensor arrays (RF MUX SQUID / TES detectors) [Kent Irwin et. al.]
  Cosmic microwave background
  High energy resolution X-ray detectors
Designed to interface to RF daughter card
  RF multiplexing filters
  Synthesizers and mixers
Status: Testing 1st prototype

## Slide 13: Rear Transition Modules
Direct connection to carrier card FPGA and to AMC cards
  3X high speed serial lanes. 2X for network and timing data stream, 1 for on-board use
  Parallel digital IO to FPGA
  Digital IO to AMC modules – can be used for clocks or interlocking
Often contain simple programmable logic (CPLD) for interlocks that need to be independent of the main FPGA code for high reliability operation.
Several variants for different applications.
  Digital IO, laser locker, low level RF, interlocks, trigger output etc.

## Slide 14: ATCA Common Platform Projects
SLAC / LCLS Low Level RF upgrade
  83 ATCA cards for SLAC LCLS operations
  Additional for SPEAR ring and FACET2 accelerator
SLAC LCLS 2, Instrumentation
  Cavity BPMs, stripline BPMs, bunch length monitor, bunch charge monitor, machine protection, beam loss monitors
  Approximately 500 ATCA cards
Cosmic microwave background telescope projects
  Approximately 100K sensor channels, 25 ATCA cards.
LCLS2 cryogenic X-ray sensor
  Approximately 10K high bandwidth sensor channels, 10 ATCA cards.
SLAC RF femtosecond timing system
  Laser locker, RF-over-fiber timing
  Approximately 30 ATCA cards