## Slide 1: R&D on μRWELL Tracking Detectors
Daojin Hong
On behalf of USTC MPGD Group
State Key Laboratory of Particle Detection and Electronics, China
University of Science and Technology of China
July 29, 2018

## Slide 2: Outline
Motivation
Overview of μRWELL detector
R&D on μRWELL detector
Summary and Outlook
R&D on the resistive layer-DLC
Detector Design
Performance study

## Slide 3: Motivation
GEM (Gas Electron Multiplier) is the solution
for SoLID tracker detector
Main requirements for GEM detectors @SoLID
  A new gaseous detector named μRWELL provides a promising candidate for SoLID tracker detector.
G. Bencivenni et al, RD51 Mini-week report
  High rate: up to the order of 1 MHz/cm2
  Position resolution: ~100μm
  Tracking efficiency: >90%
  …
μRWELL detector

## Slide 4
Overview of μRWELL detector
G. Bencivenni et al., JINST, 10, (2015), P02008
Simple and compact: no transfer or induction gap
Low mass: no stretching, no gluing and no spacers
Suitable for large area
A resistive layer for discharge suppression and current evacuation
Advantages of the μRWELL detector
Schematic of Micro Resistive Well (μRWELL) detector
Drift gap
DLC - Diamond-like carbon
resistive layer

## Slide 5: Introduction to DLC
Why use DLC as the resistive layer?
Useful link of DLC: https://en.wikipedia.org/wiki/Diamond-like_carbon
Diamond-like carbon (DLC) is a class of amorphous carbon material that displays some of the typical properties of diamond.
  Widely applied in industry
  Chemical stability
  Thermal stability
  High resistivity
Low friction factor
High hardness
Thermal conductive
…
Applications

## Slide 6: DLC deposition procedure and devices
Sample Clamping
Vacuum Pumping
Deposition
Cooling in Vacuum
Sample
taking down
Sample Support
The Magnetron sputtering system (Teer 650)
Baking base material at 70 degrees for 12 hours.
Vacuum pumping to remove the air from the chamber (should be less than 3×10-5Torr).
Start procedure to coat DLC on the pre-treated sample.
Cooling in vacuum to release the inner stress uniformly of the sample.
Sample Baking
Sample Pre-treating
Baking
DLC is deposited on a kapton substrate by the magnetron sputtering method
Main deposition procedure:
State Key Laboratory of Solid Lubrication, Lanzhou Institute of Chemical Physics, Chinese Academy of Sciences

## Slide 7: DLC Sample Performance
51 MΩ/□ | 52 MΩ/□ | 51 MΩ/□
41 MΩ/□ | 42 MΩ/□ | 42 MΩ/□
46 MΩ/□ | 45 MΩ/□ | 46 MΩ/□
A 15cm×15cm DLC Sample
Resistivity Uniformity: ~13%
Resistivity VS Time
Thickness : ～100nm
Resistivity uniformity : 13%
DLC parameters
After take down the DLC sample, we expose the DLC sample in air.
Resistivity of the DLC keep stable after 3~4 days
～65MΩ/□
Notes: Possible reason is the passivation in air

## Slide 8
Design of μRWELL Detector
Assembly
μRWELL PCB
Frame
Drift
Window
HV
Backplane for APV25 chip
2-D μRWELL Detector prototype

## Slide 9
Design of μRWELL PCB
1st
The WELL structure is split into 4 segments to avoid failure of the whole detector as a defect only in small area.
3rd

## Slide 10: Fabrication of μRWELL detector
μRWELL PCB
Fix drift electrode
Solder HV Connector
Fix μRWELL PCB
μRWELL Detector
μRWELL detector
  10cm×10cm active area
  3mm drift gap
A simple and fast assembly procedure!
(HV or signal output)

## Slide 11
Gas: Ar/CO2(70/30)
Charge collection efficiency .VS. Drift field
Collection efficiency of primary electrons .VS. Drift field
uRWELL
HV
X_ray
Pre-Amp
Spectrascopy Amplifier
MCA
~3.3 kV/cm
Gain reach the maximum @3.3kV/cm drift field
Test setup

## Slide 12
Gain of the μRWELL detector
Gain uniformity
Gain：~1800× (1±10%)
Detector area:10 cm ×10 cm
Gas: Ar/CO2=70/30
Energy resolution：~22%
Gain: could reach 104
Energy spectrum of 5.9 keV x-ray
Gain of the μRWELL detector

## Slide 13
Rate Capability
X ray counting rate is linear with current of X ray tube
Gain of the μRWELL detector drop about 2% @100kHz/cm2
P41 | P42 | P43 | P44
P31 | P32 | P33 | P34
P21 | P22 | P23 | P24
P11 | P12 | P13 | P14
X ray tube
X ray current VS counting rate
Resistive electrode
Notes: Rate capability is an important parameter for a detector in high rate occasion

## Slide 14
Scintillator for trigger
RD51 GEM Tracker
APV25 readout chip
Muon beam momentum: 150GeV/c
Beam Test Setup
SRS
S1 & S2 for trigger
GEM1
GEM2
μRWELL
GEM3
S3 & S4 for trigger
Beam line
GEM tracker
Notes: Four scintillator detectors are used as a trigger for the whole system, and three GEM detectors provide a tracker for the 2-D Urwell detector.

## Slide 15
Detection efficiency
Top layer efficiency : ~95%
Bottom layer efficiency: ~92%
Top & Bottom efficiency: ~90%
Top layer Charge/Bottom layer charge
Detection efficiency VS avalanche voltage
The Top layer induced charge is 1.9 times of Bottom layer
Parameter of readout strips
Pitch: 400 μm
Top layer: 80 μm
Bottom layer: 350 μm

## Slide 16
Position resolution
Sigma: 66 μm
Sigma: 68 μm
Notes: X position resolution(Bottom layer): 66um

## Slide 17
Micro Resistive Well (μRWELL) is a compact detector, it is simple and easy to fabricate.
Performance of a μRWELL detector is studied in lab/ muon beam.
Summary and Outlook
Outlook
Summary
R&D on larger size DLC sample.
R&D on large area μRWELL detector.
R&D on high rate μRWELL detector.
Gas gain: could reach 104
Rate capability: >100kHz/cm2
Position resolution: ～70 μm.
Detection efficiency to 150 GeV/c muon: >90%
Special Thanks to:
Teixeira, R. De Oliveira and G. Bencivenni for
providing technical support
Thanks!