## Slide 1: Progress of LHAASO-WFCTA SiPM Camera
GE, Maomao, Yunnan University
gemaomao@ynu.edu.cn
The 2nd LHAASO Collaboration Meeting in 2019 (Zhuhai)

## Slide 2: Outline
Introduction
Engineering of the WFCTA SiPM Camera
The Camera Production in 2019
A Work Plan of the Next Year
Conclusion

## Slide 3: Introduction
LHAASO-WFCTA
Functions of WFCTA
Discriminating species and measuring energy spectrum of CRs within the range from 10 TeV to 1 EeV
Phase | Configuration | Main Objective
I | Cherenkov Mode (6 WFCTs+¼WCDA) | 100 TeV-10 PeV Light components of CRs
II | Cherenkov Mode (18 WFCTs+ED+MD) | 1 PeV-100 PeV Heavy components of CRs
III | Fluorescence Mode (20 WFCTs+MD) | >100 PeV The second “knee”

## Slide 4: Introduction
The WFCTA SiPM Camera
Summary Information
A camera with an optical filter at the Mt. Haizi site
Num. of Pixel | 32×32 (1024)
F.O.V. | 16°×16°
Pixel Size | 0.5°×0.5°

## Slide 5: Introduction
Camera Stucture
Thin Film Physics
Hamamatsu
IHEP
IHEP
SCU
SCU
SCU
SCU

## Slide 6: Introduction
WFCTA SiPM Camera Review Meeting(2019 3/Jun.)

## Slide 7: Introduction
Our Tasks
Research phase (2010-early 2019) | Camera test systems research Engineering approaches research
Construction phase (early 2019-now) | Camera assembly, performances characterization, and laboratorial calibration

## Slide 8: Introduction
Camera Assembly and Test Process

## Slide 9: Engineering of the WFCTA SiPM Camera
Assembly Requirements
 | Specification | Requirement
1 | Flatness of Winston-cone entrance | ＜±0.6 mm
2 | Gap between pixels | ＜0.4 mm
3 | Gap between Sub-Clusters | ＜1 mm

## Slide 10: Engineering of the WFCTA SiPM Camera
Test Requirements
 | Characteristics | Requirements
1 | Voltage vs. Gain | 1.Uncertainty of operation voltage<5%，Statistical Uncertainty: ≤1% 2.Voltage range: 56V－61V
2 | Linearity | 1. Statistical Uncertainty: ≤1% 2. N.P.E range: 10(±5)－36000 pe
3 | Signal Resolution | 1. Statistical Uncertainty: ≤0.5% 2. N.P.E range: 10(±5)－36000 pe
4 | Temperature vs. Gain | 1. Statistical Uncertainty: ≤0.5% 2.Temperature range: 30℃－60℃ 3. Sampling proportion: 1%
5 | Baseline (pedestal) monitoring | Every electronic channel
6 | Temperature sensor monitoring | Every SiPM

## Slide 11: Engineering of the WFCTA SiPM Camera
Assembly——(1)Winston-cone Adhesion
Gluing on the supporting leg of the cones, 1 mm away from the optical surface
Fixing it in the clamping fixture

## Slide 12: Engineering of the WFCTA SiPM Camera
Assembly——(2) Assembly of SiPMs and PAB

## Slide 13: Engineering of the WFCTA SiPM Camera
Assembly——(3) Adhesion of Cones an PAB

## Slide 14: Engineering of the WFCTA SiPM Camera
Assembly——(4) Assembly of Electronics

## Slide 15: Engineering of the WFCTA SiPM Camera
Assembly——(5) SubCluster Installation

## Slide 16: Engineering of the WFCTA SiPM Camera
Assembly——(6) Wiring

## Slide 17: Engineering of the WFCTA SiPM Camera
Assembly——Flatness Measurement
Guiding rule
Caliper
 | h1 | h2 | h3 | h4 | h5 | h6 | H
Design | 25.27 | 7.50 | 7.00 | 4.00 | 169.00 | 4.10 | 216.87
Measured Tolerance |  |  |  |  |  |  | 
Specification | Requirement | Measured
Flatness of Winston-cone entrance | <±0.6mm | <±0.5mm 100 positions

## Slide 18: Engineering of the WFCTA SiPM Camera
Assembly——Gap between Pixels Measurement
Specification | Requirement | Measured
Gap between pixels | <0.4mm | <0.4mm

## Slide 19: Engineering of the WFCTA SiPM Camera
Test——Test Process

## Slide 20: Engineering of the WFCTA SiPM Camera
Test——(1) Voltage vs. Gain
Data in .txt format
All channel operation voltage for G= 8×105 Camera C06
Curves in .pdf format

## Slide 21: Engineering of the WFCTA SiPM Camera
Test——(2)Linearity
13pe
40000pe
3.5 orders of magnitude
,
~0.54%
Specification | Requirement | Measured
Stat. Uncertainty | <1% | 0.8%
N.P.E. Range | 10±5－36000 | 13－>40000

## Slide 22: Engineering of the WFCTA SiPM Camera
Test——(3)Signal Resolution
Specification | Requirement | Measured
Stat. Uncertainty | <0.5% | 0.2%
N.P.E. Range | 10±5－36000 | 13－>40000

## Slide 23: Engineering of the WFCTA SiPM Camera
Test——(4)Temperature-Gain(spot-check)
Specification | Requirement | Measured
Statistical Uncertainty | <0.5% | 0.2%
Temperature Range | 30℃－60℃ | 30℃－60℃
-1.1%/℃±0.1%/℃

## Slide 24: Engineering of the WFCTA SiPM Camera
Test——(5)Temperature-Dark Count Rate(spot-check, single channel)
Measured(qualitative)：
1. 600-800 kHz per ch. at 20℃ (0.5 pe threshold) 。
2. 10 times per 25℃。
Expected(quantitative)：
where
A is an constant,
T is the thermodynamic temperature, it is 297 K in our test.
Eg is band gap energy, crystalline silicon has a Eg of 1.1 eV at 297 K.
K is Boltzmann’s constant

## Slide 25: Engineering of the WFCTA SiPM Camera
Test——(6)Spatial Uniformity(spot-check)

## Slide 26: Engineering of the WFCTA SiPM Camera
Test——(7)Gain Monitoring(spot-check)
Pedestal
1 p.e.
2 p.e.
3 p.e.
4 p.e.
5 p.e.
6 p.e.
7 p.e.
……
The SiPM gain is given by:
Where e is the electron charge in pC.

## Slide 27: Engineering of the WFCTA SiPM Camera
Camera Test——Baseline(Pedestal)
Camera Test——Baseline Fluctuation
High Gain
Low Gain

## Slide 28: Engineering of the WFCTA SiPM Camera
Camera Test——Signal Uniformity
Camera Test——Temperature
Temperature Map of Camera C06
cooling air inlet
cooling air outlet
Ambient temperature=25℃

## Slide 29: Engineering of the WFCTA SiPM Camera
Camera Test——Signal Resolution
Camera Test——Gain Ratio

## Slide 30: Engineering of the WFCTA SiPM Camera
Camera Test——High Voltage
Camera Test——HV Deviation

## Slide 31: The Camera Production in 2019
The Overall Progress
We have delivered 5 Cameras from 16/Mar. to 25/Sep.. The production time of them are listed below：
Camera SN. | Start & Finish Date of Sub-Cluster Assembly and Test | Production Time(days) | Camera Test and Aging Time(days)
C02 | 2019/3/16-2019/4/8 | 23 | 10
C03 | 2019/5/9-2019/5/29 | 20 | 13
C04 | 2019/6/21-2019/7/7 | 16 | 6
C05 | 2019/7/23-2019/8/9 | 17 | 20
C06 | 2019/9/10-2019/9/23 | 13 | 5
30 days
22 days
16 days
31 days
32 days

## Slide 32: The Camera Production in 2019
Defective Parts
Winston-cone(14558 pieces) |  | 
Coating defects | 203 | 1.39%
No coating | 2 | 0.01%
SiPM(7136 pieces) |  | 
Soiled | 3 | 0.04%
Abnormal Baseline | 11 | 0.15%
Abnormal Noise | 13 | 0.18%
Abnormal Gain | 3 | 0.04%
Temperature sensor failed | 5 | 0.07%
AB(660 pieces) |  | 
Abnormal Baseline | 1 | 0.15%
CLB(341 pieces) |  | 
Status readback failed | 3 | 0.88%

## Slide 33: The Camera Production in 2019
Improvement
The output signal of C03 at the operation voltage given by Hamamatsu
ADC Vref of CLB
LT1963A
LDO Regulators
±1.5%
The Deviation from HVmon to HVset

## Slide 34: The Camera Production in 2019
Improvement
SiPM serial number logging error

## Slide 35: The Camera Production in 2019
Improvement
Screw thread length discrepancy

## Slide 36: The Camera Production in 2019
Improvement
The camera housing welding deformation

## Slide 37: A Work Plan of Next Year
3 Cameras (before Spring Festival)
A monochromatic (405 nm) laboratorial calibration system(Jan.)
5 Cameras (Mar.—Jun.)
6 Cameras (Aug.—Nov.)
A Multi-wavelength laboratorial calibration system (before Oct.)

## Slide 38: Conclusion
We have developed a set of engineering approaches in order to assembly and test the WFCTA SiPM camera according to the requirements.
So far, we have delivered 6 cameras including the C01 engineering prototype. These cameras operate normally at the site with well maintenance of colleagues.
Some deficiencies in the camera parts or in the assembly and test approaches have been found and eliminated.

## Slide 39
Thanks!

## Slide 40: Schematic of 1D-System

## Slide 41: Schematic of 2D-System

## Slide 42: 1D-System

## Slide 43: Darkroom and Temp-System

## Slide 44: Temperature System