## Slide 1: Design of EMuS μSR spectrometers and prototype development
Ziwen Pan, Jingyu Dong, Bangjiao Ye
State Key Laboratory of Particle Detection and Electronics, USTC
14/12/2019, 2nd workshop on CSNS Muon Source Multidisciplinary Applications

## Slide 2: Outline
Design of μSR spectrometers (Day-one and Super-μSR)
Development of the first Chinese μSR spectrometer prototype

## Slide 3: Outline
Design of μSR spectrometers (Day-one and Super-μSR)
Development of the first Chinese μSR spectrometer prototype

## Slide 4: 1 Design of muSR spectrometers
Polarized muons
Sample
Asymmetrical
decay
Positron detection
Spectrum
unfolding
Probing
Magnetism
Superconductivity
Semiconductor
Diffusion
Physical chemistry
Thin films
…
Procedure
μSR spectrum
Fit & analysis

## Slide 5: 1 Design of muSR spectrometers
Advantages
µSR: A quantum magnetism probe
Very weak effects (small moment magnetism ~ 10-3 µB/Atom)
Random magnetism (e.g. spin glasses).
Short range order (where neutron scattering is not sensitive).
Full polarization in zero field (unique measurements without disturbance of the system)
Single particle detection (with extremely high sensitivity)
No restrictions (in choice of materials to be studied)

## Slide 6
Rate
Spectrometer
Beam
Segments
Deadtime
Intensity
Frequency
a0
Spectrometer
Beam
Placements
Degrader
Time width
Polarization
1 Design of muSR spectrometers
Precision: Source

## Slide 7
Precision: Counting rate
Adding
counts
2.83 MEvents
97.97 MEvents
1 detector
48 detectors
Covering a large solid angle
Highly segmented
Short dead time to tolerate high muon beam intensity
Requirements for spectrometers:
1 Design of muSR spectrometers

## Slide 8
Precision: Asymmetry
ZF/LF
TF
a0 is the intrinsic asymmetry of a μSR spectrometer!
1 Design of muSR spectrometers

## Slide 9
As the time width of muon pulses is large, the amplitude (asymmetry) of muon spin oscillation dephase (reduces).
muons
sample
Arrangement of EMU spectrometer
1 Design of muSR spectrometers
Precision: Asymmetry
quartz

## Slide 10
muons
sample
Arrangement of EMU spectrometer
Requirements for spectrometers:
Occupying high intrinsic asymmetries (a0) in space
Optimizing detector arrangement
Optimizing degrader thickness
1 Design of muSR spectrometers
Precision: Asymmetry
quartz

## Slide 11: 1 Design of muSR spectrometers
After optimizing the intrinsic asymmetry, the oscillation becomes more significant.
Change placement
Adding degrader
Counts: 3.60105 muons
Asymmetry: 0.199
Counts: 1.79105 muons
Asymmetry: 0.376
Precision: Asymmetry

## Slide 12
ε is normalized kinetic energy of positrons
ε0 is the lower limit of primary kinetic energy of detected positrons
Asymmetrical decay
θ1
θ2
Maximum value：
Intrinsic asymmetry：
1 Design of muSR spectrometers

## Slide 13
Detector rings far from the sample occupy high asymmetry
Asymmetry can be increased by selecting positrons with high energy
Degraders can be mounted on the path of positrons penetrating detectors to block low energy positrons with low asymmetries
Detector arrangement: sphere
Amax
1 Design of muSR spectrometers

## Slide 14
Detector arrangement: cylinder
Penetration positions far from the sample occupy high asymmetries
Detector rings with small radius have high asymmetries
Asymmetry can be increased by selecting positrons with high energy
Degraders can be mounted on the path of positrons penetrating detectors to block low energy positrons with low asymmetries
1 Design of muSR spectrometers

## Slide 15
Demonstration of the theoretical calculation
1 Design of muSR spectrometers
degrader
For Z < 200 mm, the trend is similar to the theoretical calculation;
After that, asymmetry reduces as more and more positrons are not from the sample.

## Slide 16
Design of Day-one spectrometer
1 Design of muSR spectrometers
Degrader (mm) | Highest asymmetry | Average asymmetry | Event rate (Mevents/hr)
9 | 0.415 | 0.375 | 22.39
0.28 @ ISIS
0.16 @ J-PARC
256 Channels

## Slide 17
Design of Super-μSR spectrometer:
1 Design of muSR spectrometers
Cylinder
Size：101010 mm3
Number：2560
Rings：252
Solid angle：30.6%
Cone
Size ：101010 mm3
Number ：2464
Rings：132
Solid angle ：64.1%
Size ：~ 4.94.950 mm3
Number ：2686
Rings：142
Solid angle ：42.5%
Sphere

## Slide 18
Design of Super-μSR spectrometer
1 Design of muSR spectrometers
Searching high asymmetry area
Z0
Z
-Z0

## Slide 19
Moving detector ring away from the sample can have high asymmetry
The increase of asymmetry tends to saturate when moving rings
Spherical shape has the highest asymmetry
Design of Super-μSR spectrometer
1 Design of muSR spectrometers

## Slide 20
Design of Super-μSR spectrometer
1 Design of muSR spectrometers
 | Thickness (mm) | Average asymmetry | Event rate (MEvents/hr)
Cylindrical | 9 | 0.384 | 60.35
Conical | 9.5 | 0.270 | 214.9
Spherical | 9.5 | 0.409 | 18.85

## Slide 21: Outline
Design of μSR spectrometers (Day-one and Super-μSR)
Development of the first Chinese μSR spectrometer prototype

## Slide 22: 2 Development of the first Chinese μSR spectrometer prototype
Goal：constructing a 128-channel μSR spectrometer prototype
Positron detectors:
128 channels (scintillator + light guide + PMT)
Electronics：
LED / CFD + TDC
data acquisition software
Magnetic field：
zero field
longitudinal field: 0~300 Gauss
transverse field: 0~100 Gauss
Sample environment：
room temperature
sample holder

## Slide 23: 2 Development of the first Chinese μSR spectrometer prototype
What have been done:
Conceptual design and update 
update

## Slide 24: 2 Development of the first Chinese μSR spectrometer prototype
What have been done:
Detector prototype
upgrade

## Slide 25: 2 Development of the first Chinese μSR spectrometer prototype
What have been done:
Electronics and software
TDC
Software
CFD
LED
upgrade

## Slide 26
First beam test
Second beam test
Third beam test
14th Mar. 2018
25th Sept. 2018
13th Nov. 2019
2 Development of the first Chinese μSR spectrometer prototype

## Slide 27
2 Development of the first Chinese μSR spectrometer prototype

## Slide 28
2 Development of the first Chinese μSR spectrometer prototype
Main results: pulse height spectra
1st beam test
2nd / 3rd beam test

## Slide 29
2 Development of the first Chinese μSR spectrometer prototype
Main results: μSR spectra
Crosstalk
1st beam test
Unexpected oscillation
Distortion
Low counting rate
High counting rate
Crosstalk

## Slide 30
2 Development of the first Chinese μSR spectrometer prototype
Main results: μSR spectra
Low counting rate
High counting rate
2nd beam test

## Slide 31
2 Development of the first Chinese μSR spectrometer prototype
Main results: μSR spectra
Low counting rate
High counting rate
3rd beam test

## Slide 32
2 Development of the first Chinese μSR spectrometer prototype
Main results: dead time performance in comparison
2nd / 3rd beam test
After upgrading the front-end electronics (CFD → LED), distortion of μSR spectra in early times improves a lot.

## Slide 33
2 Development of the first Chinese μSR spectrometer prototype
Main results: dead time performance in comparison
3rd beam test
The performance of Chinese μSR spectrometer prototype reaches an advanced level similar to ISIS instruments.

## Slide 34
Day-one spectrometer
256 PMTs， 8 rings
Asymmetry：average-0.38, highest-0.42
Event rate：2.5107/hour
Super-μSR spectrometer
~2560 SiPMs
Asymmetry：average-0.35
Event rate：2 108/hour
Conclusion

## Slide 35
µSR: A quantum magnetism probe
Thanks!