## Slide 1: Assembly and tests of RPC at USTC for the ATLAS Phase-II Upgrade
Yuexin Ding
USTC

## Slide 2: OUTLINE
INTRODUCTION
  The ATLAS RPC model
  Current RPC systems of the ATLAS detector and the Limitations
  Upgrades of RPC detectors for HL-LHC
  BIS78 and USTC-cluster’s responsibility
SINGLET ASSEMBLY
  Gas gap training & test
  FE board soldering & connectivity test
  Assembly procedure
TEST RESULTS
  Experiment Setup
  Efficiency-HV Curves、Cluster Size、Time Resolution
CONCLUSION

## Slide 3: CONTENT
INTRODUCTION
  The ATLAS RPC model
  Current RPC systems of the ATLAS detector and the Limitations
  Upgrades of RPC detectors for HL-LHC
  BIS78 and USTC-cluster’s responsibility
SINGLET ASSEMBLY
  Gas gap training & test
  FE board soldering & connectivity test
  Assembly procedure
TEST RESULTS
  Experiment Setup
  Efficiency-HV Curves、Cluster Size、Time Resolution
CONCLUSION

## Slide 4: The ATLAS RPC model
Readout panel
Readout strip
Ground panel
Readout panel
Medium-resistive electrode
Working gas
Insulating layer
Graphite layer
Avalanche
Spacer
Key parameters of current RPC detectors | 
Time resolution | ns for gap size =
Spatial resolution | or higher with weighted charge
Rate capacity | for the designed lifetime
Efficiency | during ATLAS operation
Key parameters of current RPC detectors | 
Time resolution | 
Spatial resolution | 
Rate capacity | 
Efficiency | 
Frontend electronics
RPC structure & signal induction
Readout scheme & signal propagation
Frontend electronics
Gas gap
Eta Panel
Phi Panel

## Slide 5: Current RPC systems of the ATLAS detector
Arrangements
  One layer of RPC chamber in Barrel Outer
  Two layers of RPC chambers in Barrel Middle
  No RPCs in Barrel Inner currently
Limitations
  Geometrical acceptance
    Around 70% for BM/BO RPCs
    Due to magnetic coils and supporting structures
  Trigger efficiency
    HV (total charge delivered per count) must be lowered, to avoid accumulating charge damaging the RPC electrode due to the present RPC longevity
    The single-gap efficiency will have to be reduced on average by 15%
Notes: 97_MS_TDR P406
The motivation specially for Phase I upgrade: the barrel-endcap transition region (1.0<|𝜂|<1.3) features the highest background rates in the barrel
Motivation from P104, MuonPhaseII TDR
One unpresented motivation is the new gas mixture

## Slide 6: Upgrades of RPC detectors for HL-LHC
Upgrade method
  Installing thin-gap RPCs in the BI layer
An additional layer of BI RPCs can …
  Keep the high trigger efficiency
  Improve trigger selectivity
  Reduce trigger fake rate
  Increase geometrical coverage
Other improvements
  New front-end electronics for amplifying and discriminating the avalanche signals
Present
Eff. Accept. in worst case for
present
Notes: Motivation from P104, MuonPhaseII TDR
The motivation specially for Phase I upgrade: the barrel-endcap transition region (1.0<|𝜂|<1.3) features the highest background rates in the barrel
Motivation from P104, MuonPhaseII TDR
unpresented motivations are the new gas mixture and new trigger rate requirements

## Slide 7: BIS78 RPC
BIS78 RPC
  BIS: MDT sMDT + thin-gap RPC
  Covering the area of present BIS7 and BIS8 MDTs
  Running in present Phase-I Upgrade to verify the feasibility
New RPC design
  Gap size: 2 mm  1 mm
  Electrode thickness: 1.8 mm  1.2 mm
  Advantages
    To reduce the required voltage from 9600 V to 5400 V
    To reduce the total chamber thickness
    To improve the time resolution from about 1.1 ns to 0.4 ns and the signal-to-noise ratio

## Slide 8: USTC-cluster’s responsibility
USTC-cluster’s responsibility
  R&D for double-end readout
  Production of all BIS readout panels: PCB + paper-honeycomb
  Assembly and test of all BIS RPCs
Our finished work
  Have Assembled four RPC singlets
    BIS7s size, similar to the size of BIS small ones
    Readout panels and gas gaps rent from the Phase I upgrade project
    Front-end boards designed by Roma II with discrete elements
  Have tested the assembly RPC singlets’ performances
    Efficiency-HV curves
    Cluster Size
    Time resolution
  To show the assembly ability of USTC for the upcoming mass production

## Slide 9: CONTENT
INTRODUTION
  The ATLAS RPC model
  Current RPC systems of the ATLAS detector and the Limitations
  Upgrades of RPC detectors for HL-LHC
  BIS78 and USTC-cluster’s responsibility
SINGLET ASSEMBLY
  Gas gap training & test
  FE board soldering & connectivity test
  Assembly procedure
TEST RESULTS
  Experiment Setup
  Efficiency-HV Curves、Cluster Size、Time Resolution
CONCLUSION

## Slide 10: Gas gap training & test
Gas gap training & test
  HV (high voltage) training
  HV-LC (leakage current) test
    LC need to be lower than 3uA
Chart: High voltage [V]
| 1000 | 2000 | 3000 | 4000 | 4200 | 4400 | 4600 | 4800 | 5000 | 5200 | 5400 | 5500 | 5600 | 5700 | 5800 | 5900 | 6000 | 6100 | 6200 | 6300 | 6400 | 6500 | 6600
 | 0.0178 | 0.0388 | 0.0633 | 0.0957 | 0.105 | 0.1143 | 0.1268 | 0.1442 | 0.1754 | 0.2116 | 0.2837 | 0.3367 | 0.4134 | 0.5172 | 0.6645 | 0.8505 | 1.0792 | 1.4095 | 1.7956 | 2.292 | 3.036 | 3.931 | 5.3689
 | 0.0261 | 0.0572 | 0.0956 | 0.1455 | 0.1598 | 0.173 | 0.1892 | 0.2071 | 0.2354 | 0.2765 | 0.355 | 0.4255 | 0.5312 | 0.683 | 0.8933 | 1.1774 | 1.5773 | 2.0488 | 2.6198 | 3.4839 | 4.5401 | 6.0352 | 7.9196
 | 0.0254 | 0.056 | 0.0883 | 0.1234 | 0.1298 | 0.1347 | 0.1436 | 0.1562 | 0.172 | 0.1938 | 0.2313 | 0.2605 | 0.3054 | 0.3655 | 0.4446 | 0.5294 | 0.6609 | 0.8062 | 1.0202 | 1.3142 | 1.7313 | 2.1724 | 2.8153
 | 0.031 | 0.0652 | 0.1025 | 0.1439 | 0.1533 | 0.1628 | 0.1738 | 0.187 | 0.2023 | 0.2188 | 0.252 | 0.2784 | 0.3201 | 0.3733 | 0.4513 | 0.5651 | 0.6913 | 0.8813 | 1.155 | 1.4788 | 1.9066 | 2.5805 | 3.3184
HV-LC test result
Raw Gas Gap

## Slide 11: FE Board --- Selection of working points
Version
  a discrete-element FE board redesigned by Roma II
  One board can cover 8 readout strips
Low voltage supply （five discrete LV）
  Vpu = 1.2 V
  Vdis = 2.5 V
  VLVDS = 2.4 V
  Vamp and Vth should be scanned to select the best working point
Selection of working points
  A wide scan on Vamp and Vth
  fixed FE working point
    Phi FE boards: Vamp = 1.5 V, Vth = 3.3 V
    Eta FE boards: Vamp = 1.3 V, Vth = 3.45 V
Full Figure
Partial Figure

## Slide 12: FE board soldering
Insulating
  Tape the Mylar on the back of FE boards
Positioning
  Align the FE boards with readout strips
  Uncover the positioning holes
Soldering
  Solder six positioning holes first
  Solder the FE and the Signal/Gnd of the readout strips
  solder the FE and the LV power supply

## Slide 13: FE board connectivity test
Signal generator
  To provide 2mV pulse signal and introduce to the strips
  Connect the sync output of the signal generator to the oscilloscope as trigger.
Oscilloscope
  Observe whether a voltage jump signal which is strongly relative to the sync signal or not
    If yes, that means the FE board connect to the strips stably and is in a good condition

## Slide 14: Assembly procedure
Positioning
  Dig a hole on each readout panel
    to leave space for the HV connection on the gas gap
  Tape Forex strips along Eta and Phi strips
    To align / limit the position of the gas gap
Assembling
  Cover an additional insulating layer on the side connecting high voltage
  Use copper tape to cove the edge of the singlet
Shielding
  Connect the FE boards of one readout panel with the ground side of the other readout panel by soldering a wire

## Slide 15: Gantt chart for assembly activity
Ramping up slowly at first
  Mould making
  Material purchasing
  Warming up
2 days / singlet after ramped up

## Slide 16: Document archives
Documents has been archived
Manuals have been written down

## Slide 17: CONTENT
INTRODUTION
  The ATLAS RPC model
  Current RPC systems of the ATLAS detector and the Limitations
  Upgrades of RPC detectors for HL-LHC
  BIS78 and USTC-cluster’s responsibility
SINGLET ASSEMBLY
  FE board and Soldering
  Assembly Procedure
  Assembly Checkout
TEST RESULTS
  Experiment Setup
  Efficiency-HV Curves、Cluster Size、Time Resolution
CONCLUSION

## Slide 18: Experiment Setup
RPC singlet under tests
  4 BIS7s RPC singlets, for efficiency-HV scan
Trigger system
  The coincidence of two scintillators for trigger
    One above RPCs, the other below RPCs
    Along the long side covers 4 FE boards , along the short side covers 1 FE board
Data-taking system
  TDC: V1190, 128-ch
  Connecting with FE boards
cosmic ray
Simple structure of trigger system
scintillator
RPC singlet

## Slide 19: Experiment Setup
Phi FE #1
Phi FE #2
Phi FE #8
Eta FE #1
Eta FE #2
Eta FE #3
Eta FE #4
…….
Phi FE #3
Phi FE #4
Eta panel below the gap, facing positive HV side
Phi panel above the gap, facing the GND side
Scintillator (for trigger)
0.8m * 0.2m Covering 32 channels
RPC Singlet (for testing)
1.8m * 1m large area
64 (Phi) + 32 (Eta) channels

## Slide 20: Process of signal transmission and data taking
RPC
  Avalanche signal generate in the gas gap
  Signals propagate along the readout strips
FE boards
  Current signal  (matching resistor)  voltage signal
  Amplification
  Discrimination
TDC
  To record the time of rising and trailing edges
  To convert digital signals to data
PC
  To control TDC
  Data storage onto disks
  Data processing
RPC (resistor)  FE
FE  (FFCLVDS)  TDC

## Slide 21: Efficiency-HV Curves
Criteria: Efficiency > 95% HV 6200V

## Slide 22: Cluster Size
Definition:
  The number of strips fired by the same incoming muon
Criteria: average cluster size < 3
Phi panel
Eta panel

## Slide 23: Time Resolution
HV = 6200 V
Including from TDC, cables and FE boards
Phi
Eta

## Slide 24: CONTENT
INTRODUTION
  The ATLAS RPC model
  Current RPC systems of the ATLAS detector and the Limitations
  Upgrades of RPC detectors for HL-LHC
  BIS78 and USTC-cluster’s responsibility
SINGLET ASSEMBLY
  Gas gap training & test
  FE board soldering & connectivity test
  Assembly Procedure
TEST RESULTS
  Experiment Setup
  Efficiency-HV Curves、Cluster Size、Time Resolution
CONCLUSION

## Slide 25: Conclusion
Four singlets assembled and tested
  DAQ set up
  Manuals written down
  Data stored
  Criteria reached
We have the assembly ability for the upcoming mass production

## Slide 26: Thanks!