## Slide 1: IHEP-EDHP-AMC a MicroTCA.4 Based Digital Processing AMC Module
Long Wei
Project group of domestic MTCA.4 platform
MTCA Workshop, Beijing 2021.8
China Spallation Neutron Source

## Slide 2: IHEP-EDHP-AMC Overview
CSNS & CSNS II
SAPS-TP
SAPS
Will be built
Dongguan City, Guangdong Province, China
Introduction to our facilities.

## Slide 3: IHEP-EDHP-AMC Overview
CSNS (China Spallation Neutron Source) – China’s first pulsed neutron source.
80 MeV H- Linac
1.6 GeV RCS
Target station
Phase | CSNS I | CSNS II
Beam power on target | 100 kW | 500 kW
Linac energy | 80MeV | 300MeV
Extraction beam energy | 1.6GeV | 1.6GeV
Average Beam current | 62.5µA | 312.5µA
Repetition | 25Hz | 25Hz
Protons per pulse | 1.56E13 | 7.8E13
CSNS II upgrade

## Slide 4: IHEP-EDHP-AMC Overview
SAPS Southern Advanced Photon Source
4th generation medium energy synchrotron radiation source
SAPS-TP Southern Advanced Photon Source Test platform
Storage ring Beam energy | 3.5GeV
emittance | 31.7pm·rad
circumference | 1080m
RF frequency | 166.6MHz
SAPS
SAPS
SAPS-TP
SAPS-TP
Superconducting RF Hall
optical experiment Hall
low temperature Hall
High-accuracy measurement Hall
Comprehensive laboratory

## Slide 5: IHEP-EDHP-AMC Overview
Multifunction digital processing AMC Module
Standard universal MTCA.4 Embedded digital hardware platform
CSNS II and SAPS
  RF system & Control system & Beam diagnostics system
SAPS-TP
  superconducting cavity Vertical Test / Horizontal Test
Other applications
Different types of RTMs
FMCs
Chassis
Setup
implement
MCH
A domestic multifunction digital processing AMC module is developed.

## Slide 6: functionality
MTCA.4 standard & μRTM ZONE3 Class A1.1CO
Double-width mid-size
8 Channels 125 MSPS 16-bit ADC
2 Channels 500 MSPS 16-bit DAC
XC7Z045 Zynq-7000 SoC
2GBytes DDR3 memory size
1 FMC LPC Slot (VITA 57.1)
A 4 lanes QSFP cage
Front panel digital I/O (10 pins LVTTL)
A Gigabit Ethernet interface
Front panel Clock input
A front-panel USB provides console DisplayPort
FMC
CLK
ETH
DIO
QSFP
USB
SD

## Slide 7: IHEP-EDHP-AMC Block Diagram & Overall Design
ZYNQ SOC
ARM Cortex-A9
Kintex-7
FPGA
AXI GP&HP bus
Ext Clock, SMA
FMC LPC
USB
JTAG
LED
Digital I/O
QSFP
E-Net RJ45
Dual ADC
Dual ADC
Dual ADC
Dual ADC
Clock Distribution
Dual DAC
MMC Connector
AMC Connector
Zone3 Connector
Hot-swap SW
EPICS IOC
LLRF
Algorithm
FMC Board
Motor Controller
Piezo Controller
optical fiber communication
…
RTM Module
Down-Converter
Up-Converter
Direct Sampling
…
A goal: one AMC board can constitute a complete system and can implement complete project applications.
Network communication
Remote monitoring
OPI
…
Trigger Signals
Inter-Lock Signals
…
Timing & Interlock
Timestamp…
Ext Clock source

## Slide 8: IHEP-EDHP-AMC Block Diagram & Overall Design
SAPS-TP 648MHz Elliptical Cavities Single Cavity Regulation based on MicroTCA.4
Piezo Driver
MTCA.4 AMC
MTCA.4 RTM
Down & UP
Converter
ZYNQ
FPGA
ARM
16bits DAC 2Ch
LP
Level conv.
LP
Motor Driver
FMC
5V pulse signal
0~10V
Cavity detuning control
LLRF system
Only one IHEP-EDHP-AMC module + 1 RTM + 1 FMC

## Slide 9: IHEP-EDHP-AMC MMC
MMC Controller
Based on N.A.T NAMC-MMC-RefDesign
Developed MMC controller mezzanine Card
MMC Controller in AMC
MMC Controller Test
Module operational state management
Sensor Reading Test
Temperature threshold event Test
82℃
Fan speed increase
30℃
Fan speed decrease
Hot swap management works well.
Test environment:
Commercial nVent MTCA.4 chassis
N.A.T NAT-MCH-PHYS80 MCH card
3 phases of development:
HW SW based on N.A.T MMC RefDesign
Develop our own Software
Design our own MMC (HW SW)

## Slide 10: IHEP-EDHP-AMC AD/DA
AD: Analog Devices AD9268
  125MSPS 16bit
  Bandwidth: 650MHz
  Output Interface: DDR LVDS
Compatible analog front-end circuit design
  Intermediate frequency range: 5-100MHz
  Direct Sampling: 5-400MHz (650MHz? Try)
  DC coupled
DA: Analog Devices AD9783
  500MSPS 16bit
  Output Interface: DDR LVDS
2 channels DA 8 channels AD
DC/AC coupled is optional

## Slide 11: IHEP-EDHP-AMC AD/DA
AD Test
  IHEP-EDHP-AMC + SIS8900 RTM
  Signal Source: 20MHz 3dBm
  125MSPS
SNR: 75.9~76.2dBFS > 74dBFS
SFDR: 86.2~92.1dBFS > 75dBFS
8 Chs FFT
FFT

## Slide 12: IHEP-EDHP-AMC AD/DA
Channel isolation
 | AD1_CH0 | AD1_CH1 | AD2_CH0 | AD2_CH1 | AD3_CH0 | AD3_CH1 | AD4_CH0 | AD4_CH1
AD1_CH0 |  | 74.95 |  |  |  |  |  | 
AD1_CH1 | 84.32 |  | 80.89 |  |  |  |  | 
AD2_CH0 |  | 75.22 |  | 77.1 |  |  |  | 
AD2_CH1 |  |  | 82.16 |  | 80.82 |  |  | 
AD3_CH0 |  |  |  | 87.3 |  | 79.46 |  | 
AD3_CH1 |  |  |  |  | 78.89 |  | 85.59 | 
AD4_CH0 |  |  |  |  |  | 94.8 |  | 85.13
AD4_CH1 |  |  |  |  |  |  | 78.1 | 
Isolation of adjacent channels
Isolation Test
DA Spectrum
DA output signal: 10MHz
SNR: 78.1 dB
some larger harmonic component
  + Adaptor RTM

## Slide 13: IHEP-EDHP-AMC Clock Distribution
ultra-low clock jitter
  648MHz direct sampling LLRF 0.1° jitter (<423fs pp)
  Jitter attenuator + Low phase noise clock generator
  Theoretical clock jitter ≈ 240fs pp (170fs RMS)
IN0
IN1
IN3
Jitter
attenuator
IN3
Front panel
Clock input
backplane
TCLKA
RTM CLK0
Quartz
Backplane
TCLKB
RTM CLK1-4
Clock
generator
ADC1
ADC2
ADC3
ADC4
Clock MUX

## Slide 14: IHEP-EDHP-AMC FMC
Compatible with FMC LPC standard & 4 high-speed serial transceiver pairs.
FMC LPC引脚
Compatible commercial FMC cards
DFMC-MD22
DFMC-AD16
DFMC-UNI-IO
DFMC-SFP4

## Slide 15: IHEP-EDHP-AMC Software
Embedded Linux Development for ZYNQ chip
ZYNQ
ARM
Embedded Linux
EPICS IOC
Open source Linux kernel
Tool: Xilinx petalinux
Linux Boot Image
Linux Kernel Image
23MB
45MB
FPGA
AXI DMA
DMA Proxy Driver
EPICS IOC Demon
PV: AI AO Waveform

## Slide 16: IHEP-EDHP-AMC Software
Embedded Linux Development for ZYNQ chip
User data
FIFO
AXI Master
AXI DMA IP

## Slide 17: Others
The domestic Down and up converter RTM and Direct sampling RTM is developed by the colleagues in the project group.
  Based on the requirement of SAPS and CSNS II.
Provided by: Xinpeng Ma and nan Gan
The details will be shown in a later report by Dr. Ma
“MicroTCA.4 in LLRF of BEPCII and HEPS”

## Slide 18: Others
The domestic MTCA.4 chassis is developed
Cooperation manufacturer - Yzitech company.
Provided by: Yzitech
The details will be shown in a later report by Yzitech company
“MicroTCA system engineering, developing and domestic production”
Demostic MTCA.4 chassis
Demostic MTCA.4 chassis test

## Slide 19: Thanks for your attention!