## Slide 1: Preliminary Considerations on the beam test
F. Grancagnolo
Nov. 25, 2021

## Slide 2: The test
Claudio CAPUTO UC Louvain
Federica CUNA INFN Lecce
Nicola DE FILIPPIS INFN Bari
Francesco GRANCAGNOLO INFN Lecce
Matteo GRECO INFN Lecce
Sergey GRIBANOV BINP Novosibirsk
Kurtis JOHNSON U of Florida
Sasha POPOV BINP Novosibirsk
Angela TALIERCIO UC Louvain
Shuiting XIN IHEP Beijing
Purpose
Demonstrate the ability to count clusters at a fixed βγ (e.g. muons at a fixed momentum – 165 GeV/c) by changing: - the cell size (1 – 3 cm) - the track angle (0° to 60°) - the gas mixture (90/10: 12 cl/cm, 80/20: 20 cl/cm, for m.i.p.)
Establish the limiting parameters for an efficient cluster counting: - cluster density as a function of impact parameter - space charge (by changing gas gain, sense wire diameter, track angle) - gas gain stability
Train different cluster counting algorithms
SHIFTERS

## Slide 3: The setup
HV cable
inlet gas
manifold
outlet gas
manifold
HV distr.
box
DRS
signal
box
signal cables
(40 cm)
upstream
trigger tile
downstream
trigger tile

## Slide 4: Drift tubes schematics
Tube positioning reshuffled during the test
Also, 1 cm tubes are 40 cm long
red 10 μm wire (Mo)
blue 15 μm wire (Mo)
green 20 μm wire (W)
orange 25 μm wire (W)
violet 40 μm wire (Al)

## Slide 5: The Hardware
The portable gas system
The HV system
The gain 10 - 1.7 GHz amplifier (Phillips 775)
The rotating table
The drift tubes
The gas controllers

## Slide 6: The DAQ board: DRS4
16 channels data acquisition board
designed and used by the
MEG2 experiment at PSI
(μ  e + γ)
see next presentation by G. Chiarello

## Slide 7: Cosmic rays display

## Slide 8: Analysis plan
Determine gas gain for the different (9) configurations and for different gas mixtures
Study the space charge limitations for an efficient cluster collection
Train different cluster counting algorithms by comparing to a well known method (PeakFit)
Check Poisson nature of cluster counting (tube size, track angle, gas mixture)

## Slide 9: Some considerations on gas gain
Capacitance per unit length (cylinder approximation):
CL = 2πε0/ln(R/rw) Ctube = CL × L
Inductance per unit length (cylinder approximation):
LL = 2×10-7 ln(R/rw) μH/m
our drift tubes: L = 30, 40 cm R = 0.4, 0.9, 1.4 cm
  rw = 5.0, 7.5, 10.0, 12.5, 20.0 μm
Characteristic Impedance: Z = √L/C
at gas gain = 105, one single electron deposits a charge
Q0 = 1.6×10-19×105 Coul = 16 fCoul
and the pulse height generated by a single electron is:
  ΔV = Q0 /Ctube

## Slide 10: Charge distribution
equivalent circuit
HV
1 MΩ
10 MΩ
1 nF
1 nF
330 Ω
+
Z
50 Ω
Impedance mismatch: ΔVreflect = (Z−330)/(330+Z) × ΔV ΔVtransmit = 660/(330+Z) × ΔV
Current divider: ΔV' = 330/(330+50) × ΔVtransmit = 0.87 × ΔVtransmit
combined result: ΔV' = 330/(330+50) × 660/(330+Z) × Q0/Ctube

## Slide 11: A few numbers
configuration number | R | 2×rw | L | CL | LL | Ctube | Z | ΔV (105 gain) | ΔV' (105 gain)
1 | 0.4 cm | 10 μm | 0.4 m | 8.37 pF/m | 1.34 μH/m | 3.35pF | 400 Ω | 4.78 mV | 3.75 mV
3 | 0.4 cm | 15 μm | 0.4 m | 8.91 pF/m | 1.26 μH/m | 3.56 pF | 376 Ω | 4.49 mV | 3.65 mV
5 - 7 | 0.4 cm | 20 μm | 0.4 m | 9.33 pF/m | 1.20 μH/m | 3.73 pF | 359 Ω | 4.29 mV | 3.57 mV
9 - 11 | 0.4 cm | 25 μm | 0.4 m | 9.69 pF/m | 1.15 μH/m | 3.88 pF | 344 Ω | 4.12 mV | 3.50 mV
13 | 0.9 cm | 20 μm | 0.3 m | 8.22 pF/m | 1.36 μH/m | 2.47 pF | 407 Ω | 6.48 mV | 5.04 mV
15 | 0.9 cm | 25 μm | 0.3 m | 8.50 pF/m | 1.32 μH/m | 2.55 pF | 394 Ω | 6.27 mV | 4.96 mV
17 | 0.9 cm | 40 μm | 0.3 m | 9.15 pF/m | 1.22 μH/m | 2.75 pF | 365 Ω | 5.82 mV | 4.80 mV
19 | 1.4 cm | 25 μm | 0.3 m | 7.96 pF/m | 0.94 μH/m | 2.39 pF | 344 Ω | 6.69 mV | 5.02 mV
21 | 1.4 cm | 40 μm | 0.3 m | 8.54 pF/m | 0.85 μH/m | 2.56 pF | 315 Ω | 6.25 mV | 5.30 mV
ΔV (105 gain) = single electron pulse height for an HV producing a
reduced electric field: E/P, corresponding to a gas gain of 105

## Slide 12: Determining HV (105 gain)
Garfield run on different
configurations by Shuiting
for P = 760 torr, T = 20°C
|
renormalized for
P = 725 torr and extrapolated for
10 and 40 μm wires
|
and readjusted by
observing average
maximum pulse height

## Slide 13: HV table (He/iC4H10 = 90/10, P = 725 torr)
configuration number | HV | relative gas gain (ΔV)
1 | 1230 V | 2.6
3 | 1255 V | 1.4
5 - 7 | 1305 V | 1.3
9 - 11 | 1330 V | 0. ...
13 | 1470 V | 0.6
15 | 1545 V | 0.9
17 | 1670 V | 0. ...
19 | 1620 V | 0. ...
21 | 1765 V | 0. ...
1st attempt (from Garfield)

## Slide 14: HV table (He/iC4H10 = 90/10, P = 725 torr)
configuration number | HV | relative gas gain (ΔV)
1 | 1230 V | 2.6
3 | 1255 V | 1.4
5 - 7 | 1305 V | 1.3
9 - 11 | 1330 V | 0. ...
13 | 1470 V | 0.6
15 | 1545 V | 0.9
17 | 1670 V | 0. ...
19 | 1620 V | 0. ...
21 | 1765 V | 0. ...
HV | relative gas gain (ΔV) | actual gas gain (ΔV')
1230 V – 30 V | 1.2 | 1.47
1255 V – 10 V | 1.1 | 1.29
1305 V – 5 V | 1.1 | 1.31
1330 V + 10 V | 1.1 | 1.27
1470 V + 25 V | 1.1 | 1.57
1545 V + 5 V | 1.0 | 1.49
1670 V + 50 V | 1.2 | 1.92
1620 V + 50 V | 1.1 | 2.00
1765 V + 45 V | 1.0 | 1.73
1st attempt (from Garfield)
2nd attempt (corrections to Garfield)

## Slide 15: Average maximum pulse height
Assume that the average maximum pulse height ≅
single electron pulse height
and fit to a Landau distribution by taking into account the amplifier gain (see C. Caputo talk)
amplifier gain =10
maximum pulse height

## Slide 16
red 10 μm wire (Mo)
1 cm tube blue 15 μm wire (Mo)
2 cm tube green 20 μm wire (W)
3 cm tube orange 25 μm wire (W)
violet 40 μm wire (Al)
P = 725±5 Torr
He/iC4H10 = 90/10
5
4
3
2