## Slide 1: Beam position monitor preliminary design for the CEPC
Jun He, Yanfeng Sui
Contributions from many colleagures
2022-10-28

## Slide 2
Outline
Beam Position Monitors overview
1
CEPC BPM design considering
2
Some progress on BPMs study
3
Special BPMs
4
Lessons from TDR comments
5
Summary
6

## Slide 3: Overview: Why BPMs important
BPMs are the most important system of instrumentation. Orbit measurement (bunch by bunch, turn by turn), orbit correction, bunch feedback, bunch current intensity measurement, optic functions measurement (beta-function, dispersion, tune, coupling, chromaticity) RF frequency measurement, beam loss (trip) diagnosis are depending on it.
Beam
(x,y)
A
B
C
D
Bunch charge ∝∑Vi
;
Schematic of a BPM
Providing the information of position(x,y)

## Slide 4: BPM system in 1 minute
3 parts: Detector + cable + readout electronics
A long coaxial cable
Detector
beam
~40% of the cost
BPM signal
Filter +ADC
Electronics
Cable: LMR 195 bundled
Electronics developed by BI group
~50% of the cost
Detector with four buttons

## Slide 5: BPMs number and distribution
LINAC: 140
DR: 40
Booster: 2400
~1800
~1800
Collider: ~3600
 | Numbers of BPM | Machine Size
BEPCII | 165(19+15+66×2) | 240 m
HEPS | ~700(40+80+590) | 1.4 km
CEPC | ~6200 | 100 km
FCC-ee[1] | ~6000 | ~92 km
Cost：about ¥100,000 / BPM (base on home-made electronic )
Totol length of ESBS+FAS+PSPAS+SAS+TAS: 1800 m

## Slide 6: BPMs overview
The total number of the BPMs has been increased compared with CDR because of the changing of lattice.
 | LINAC | Damping ring | Booster | e+ ring | e- ring | Transfer line | Total
CDR Button | 1 | 40 | 1808 | 1450 | 1450 | / | ~4889
CDR Strip | 140 | / | / | / | / | / | 
TDR Button | 1 | 40 | 2408 | 1772 | 1772 | / | ~6133
TDR Strip | 140 | / | / | / | / | / | 
Design principles:
LINAC – stripline type , Ring – button-type
4 BPMs /2π phase advanced (363 β oscillation period →445Htt)
1 BPM /10-20 m in LINAC 4. Electronics in the tunnel …

## Slide 7
Beam Position Monitors overview
CEPC BPM design considering
Some progress on BPMs study
Special BPMs
Lessons from TDR comments

## Slide 8: CEPC BPM design: stripline-type
Schematic of stripline BPM: (a) Front view; (b) Side view
A picture of a stripline BPM(external, interior Views)

## Slide 9: CEPC BPM design: stripline-type
Basic Requirements: The mechanical parameters (α,t,h,rin) of the strip should make the characteristic impedance of the strip Z0= 50 Ω
E-field of a stripline in different eigenmodes, (a) Sum (b) Quadrupole(c) Dipole (d) Single electrode
Higher requirements

## Slide 10: Strip impedance design
CST simulation and a time-domain reflectometer (TDR) results*
(a–b) impedance for different modes vs. h and α by CST; (c) measured by TDR(Tektronix DSA8200 ). Experimental results are in a good agreement with simulations (below 3%).
*Jun He et,al.Measurement Science and Technology 33 (2022) 115106, https://doi.org/10.1088/1361-6501/ac8277
(c)
Z vs h
Z vs α

## Slide 11: Stripline design considering
, N=0, 1, 2
rbeam is the radius of the beam stay-clear area, 1.2 mm is the minimum thickness of the strip when the longitudinal length is greater than 100 mm (enough mechanical strength), A1 and A2 are the mechanical size limits for the flange and the commercial feedthroughs, respectively.
Location | rin | t | h | α | θ | Zsingle | Lstrip | ZCST
Unit | mm | mm | mm | degree | degree | Ω | mm | Ω
LINAC(FAS+SAS) | 15 | 1.5 | 2.2 | 30 | 0 | 50.8 | 150 | 50.3±0.5
LINAC(TAS) | 10 | 1.2 | 0.6 | 30 | 0 | 49.9 | 150 | 50.4±0.5
Mechanical parameters of the CEPC stripline BPM
A button-type also could be used for LINAC, structure is simpler but signals are smaller than strip.
After considering all the associated factors mentioned above, the parameters are determined.

## Slide 12: CEPC BPM design: Button-type
U=Zt×Ibeam
The BPM design always is in a dilemma: the bigger transfer impedance (Zt) means a better resolution but also a more serious beam instability
Several tips to depress the wakefield without decrease signal too much:
1. Impedance optimization (to reduce the reflection)
2. Smaller gap (avoiding short)
3. To hide the sealed dielectric, Bell- shape (SIRIUS) or bigger tb
4. A design without skirt
5. A design of asymmetric structure
6. Dielectric with a smaller εr (To replace ceramics by glass )
Ibeam

## Slide 13: Button BPM design : Rb, tb, gb
Three important parameters for button-type BPMs: radius rb, height tb and gap gb.
Mechanical drawing and X-ray tomography of a button pickup

## Slide 14: CEPC BPM design: Button-type
Design goal: To increase the signal and decrease the wakefield.
Increase | Feature | Advantage | Disadvantage
rb | Larger area and Cb | Higher signal level | Lower resolution
tb | Larger volume | Lower wake impedance | Lower resolution
gb | More trapped modes | Higher resolution | Higher wake impedance
Mechanical structure of a BPM and a electrode

## Slide 15: CEPC BPM design: Button-type
Lab | Button radius | Half Aperture | Button Capacitance | Transfer impedance | Power
 | rb (mm) | b (mm) | Cb (pF) | Zt (Ω) | (dBm)
SPring8 | 9 | 15 | 2.5* | 1.414 | -1.6
APS-U | 5 | 11 | 2.6 | 0.595 | -9.2
ESRF | 5.5 | 15 | 2.5* | 0.528 | -10.2
BEPCII | 7.5 | 30 | 2.2 | 0.491 | -10.7
SSRF-Arc | 5 | 15.3 | 2.4 | 0.428 | -11.9
DIMOND | 5.4 | 18.5 | 2.5* | 0.413 | -12.3
HEPS | 4 | 11 | 2.4 | 0.381 | -12.9
APS | 5 | 19 | 2.6 | 0.344 | -13.9
SPring8-II | 3.5 | 10 | 2.5* | 0.321 | -14.5
ESRF-EBS1 | 4 | 13.3 | 2.5* | 0.315 | -14.6
TPS | 4 | 18 | 3 | 0.233 | -17.5
ALBA | 3.5 | 14 | 2.7 | 0.229 | -17.5
SIRIUS | 3 | 11 | 2.3 | 0.214 | -17.9
ILSF | 3.5 | 15 | 2.5 | 0.214 | -18.0
NSLS-II | 3.5 | 20.5 | 2 | 0.156 | -20.5
MAX-IV | 2 | 11 | 0.6 | 0.095 | -24.4
SKEKB | 3 | 45 | 2.5 | 0.052 | -30.2
CEPC | ? | 28 | ？ | ? | ?
Transfer impedance and signal power (I= 200 mA) for light sources and colliders[5]

## Slide 16
Responding
Mechanical parameters
Different beam parameters
Noise Figure
Cable attenuation
Sensitivity coefficient
Changing beam position
Different Bandwidth：Bunch by Bunch, Turn by Turn…
Resolution
Signal-to-noise ratio
Button BPM design flow

## Slide 17: CEPC BPM design: Button-type
Parameters of the CEPC button-type BPM for different mode
 | Value
Rb | 3 mm
tb | 4 mm
gb | 0.25 mm
Cb | 2.8 pF
The time structure of a bunch in different mode
Preliminary design parameters

## Slide 18: Accuracy Precision Resolution Repeatability
Good resolution
Bad resolution
Good accuracy
Resolution for BPM ~ Precision ~ Repeatability in other measurement
Bad accuracy
Good resolution
Bad
Good Repeatability
BPM measurement
Usually measurement
Bad Repeatability
Definition of measurement terms

## Slide 19: Accuracy & Resolution
Resolution: the ability for measuring small displacement variations. In contrast to the accuracy, relative values are compared here. Resolution is much better than the accuracy. It depends strongly on the measurement time, averaging can exceed the accuracy by a factor of 100.
Accuracy: the ability of position reading relative to a mechanical fix-point or to any other absolutely known axis, e.g., the symmetry axis of a quadrupole, mainly influenced by the BPM's mechanical tolerances as well as the long-term stability of the mechanical alignment.
Better resolution
Better accuracy

## Slide 20: Accuracy & Resolution
 | BEPCII (800/200μm) | HEPS (10/1 μm) | CEPC
Resolution | ＜ 10 μm （2s） | 0.1 μm (45 μs) | ~1 μm (4 s)
Accurancy | 100 μm | ~30 μm | ~50 μm
;
x=A0+A1Xraw+A2Yraw
Method: least square method
A baseline

## Slide 21: Accurancy in a BPM pipe
＜500 μm
＜300 μm
＜100 μm
＜100 μm
＜100 μm
＜50 μm
Higher order ~ smaller error ~more Ai ~more computing resource
n=1
n=1
n=1
n=3
n=5
n=7

## Slide 22: Symmetry of four electrodes
Although the offset would be eliminate after BBA, a correction for the coupling error is difficult; in fact, it can only be reduced if it is measured and compensated for.
Difference of four electrodes will bring additional error.

## Slide 23: BPM resolution of BEPCII
Resolution
Unit: mm
BEPCII BPM resolution
The resolution of BEPCII are on the level of several μm.

## Slide 24
Beam Position Monitors overview
CEPC BPM design considering
Some progress on BPMs study
Special BPMs
Lessons from TDR comments

## Slide 25: Study on button pickup
The structure and technology of button pick-up have been studied[2-3].
Better impedance matching
Smaller wakefield

## Slide 26: Study on home-made electronics
The home-made BPM readout electronics are about the half cost and the same performance (resolution <0.1 μm) compared with the commercial products*.
*Yanfeng Sui, CEPC Instrumentation development status, this workshop.
The picture of electronics made by BI Group
90 sets of electronics running online now(BEPCII) 700 sets will be used in the HEPS

## Slide 27: Study on mechanical tolerance
The effect of button positioning accuracy and BPM pipe mechanical tolerance have been studied[5].

## Slide 28
Beam Position Monitors overview
CEPC BPM design considering
Some progress on BPMs study
Special BPMs (near the IP )
Lessons from TDR comments

## Slide 29: Nearest BPM → beam position of IP
Beam position of IP was inferred by the position at nearest BPM. Blue & Red 4-electrode BPM. Black- special BPM(8-electrode in KEKB and BEPCII, strip in CERN)
?
?
?

## Slide 30: IP BPM: off-centre beam
4-electrode BPM
BEPCII 8-electrode BPM
The filter in the BPM readout electronics will bring a long tail which make the detection become difficult.

## Slide 31: KEKB & LHC special BPM
Embedded collimator BPMs in CERN
M. Arinaga et al. Nuclear Instruments and Methods in Physics Research A 499 (2003) 100–137
Gianluca Valentino et al. Physical review accelerators and beams 20, 081002 (2017)
BPMs in KEKB next to IP

## Slide 32: KEKB & LHC special BPM
KEK Makoto Tobiyama：The 8-bpm monitor installed in the KEKB has been working, in some degree—we have used them mainly as the nearest bpm to the IP, such as to check the collision timing, to get the beam orbit with single beam condition, or used through fast beam switch to separate the bpm signal from HER and LER. I remember it was not so stable, not so accurate as the normal position monitor-- so we could not use them as IP feedback control. In SuperKEKB, as the crossing angle is much larger than KEKB, the nearest BPMs to the IP (QC1) are placed after the separation of the two beams and have independent bpm chamber with four buttons. （update on 2022)
CERN Rhodri Jones：At the LHC we use directional stripline couplers to distinguish between the two beams. However, the isolation between beams is not perfect and we also have to rely on temporal separation for the best results. Ways to circumvent this problem are currently being studied for the high luminosity LHC upgrade. Depending on the relative intensities another option might be to rely the sign of the signal and use button BPMs with diode detection. （ update 2022）

## Slide 33
Beam Position Monitors overview
CEPC BPM design considering
Some progress on BPMs study
Special BPMs (next to the IP )
Lessons from TDR comments

## Slide 34: Review’s comments for the BPMs CDR
1. Depending on how close the closest BPMs to the interaction points will be located (0.5 m was mentioned), consider use of directional BPMs there. consider use of fancier BPMs at IPs also (e.g. OCTPOS at KEKB) √
2. modern standards include built-in calibration capabilities √
3. consider use of compensated diode detection ( Gasior at CERN) for IP BPMs √
4. consider adding quadrupole shunts to determine difference between electrical center of BPM and mechanical center of quadrupoles, see pages 73-74 of book “Measurement and Control of Charged Particle Beams” for nice examples (will be done on the HEPS at first)
5. feedthrough material considerations if high temperature baking to be performed √
6.as mentioned during HF2014 discussion by Manfred Wendt : don’t lock in ADC sampling frequency now as technology is advancing rapidly √
7. strong encouragement to provide a readout method that is independent of timing from control system (Hermann Schmickler), referring to a lessons-learned from LEP commissioning; i.e. a self-trigger mode (need to be study in future)
8. as usual: radiation aspects with electronics in tunnel – see the presentation by Luigi Salvatore Esposito on “Shielding of Electronics in the Tunnel” from HF2014 √
9. synchrotron radiation causing movement of the vacuum chambers and devices due to thermal loading seems unavoidable, KEKB has added capacitive sensors to measure the relative position of the BPMs and adjacent sextupoles,temperature sensors were very useful for beam steering in PEP2 √

## Slide 35
2: modern standards include built-in calibration capabilities
Pilot tone box for calibration will be used on the HEPS
A picture taken on yesterday!

## Slide 36
3. consider use of compensated diode detection ( Gasior at CERN) for IP BPMs
A similar system design by Gasior which has been realized on BEPCII

## Slide 37
5. feedthrough material considerations if high temperature baking to be performed.
Baking at 200℃

## Slide 38
Effect of the temperature: the radiation have less impact than the environment temperature change base on the results of BEPCII.
Ground
Displacement sensor
Temperature sensor
beam
Experiment on BII
9. synchrotron radiation causing movement of the vacuum chambers and devices due to thermal loading seems unavoidable…

## Slide 39: Effects of temperature
BPM reading Δ40μm ~ displacement sensor reading Δ 15 μm ~ Temperature 0.5 ℃
Shut down
Re-injection
The temperature difference with /without beam is 2 ℃, the difference is much smaller for different current, the position difference is 15 μm in a run (150~250 mA)

## Slide 40: Summary
The BPM system design are finished but only with some preliminary results, more details should be added. BPM requirements (resolution and accurancy) need to be further determined.
The BPM next to the IP should be studied and designed carefully.
Different project (BEPCII → HEPS → CEPC) help to deepen the research of BPM system. The resolution and accuracy of BPMs are improving with the progress of technology.

## Slide 41: Summary
I couldn’t agree more about following conclusions given by Manfred Wendt (CERN).

## Slide 42
Thank you for your attention! 
Some pubulications :
[1] S. J. Wei, Q. Ye, Y. Y. Du et al., Brief Introduction to HEPS BPM Electronics Development, in Proceedings of IBIC2020, Brazil, (2020).
[2] J. He, Y. F. Sui, Y. Li, et al., Beam position monitor design for the High Energy Photon Source. Meas. Sci. Technol. 33, 115106 (2022). https://doi.org/10.1088/1361-6501/ac8277
[3] Y. Li, J. He, Y. F. Sui, et al., Development of BPM feedthroughs for the High Energy Photon Source, Radiat Detect. Technol. Methods. (2022). https://doi.org/10.1007/s41605-022-00345-1
[4] JunHe,YanfengSui, JunhuiYue, JiaosheCao et al., Design and optimization of a Goubau line for calibration of BPMs for particle accelerators, Nucl. Instrum. Methods Phys. Res. A (2022), https://doi.org/10.1016/j.nima.2022.167635
[5] JunHe,YanfengSui, JunhuiYue, JiaosheCao et al., Design and fabrication of button-style beam position monitors for the HEPS synchrotron light facility, Nucl. Sci. Tech (2022) Accept.

## Slide 43
Back up

## Slide 44: Reading error due to 4-electrode difference
ME will be controlled by BPM manufacturing process,QM controlled by alignment, QE is the real important parameters for accelerator, it can be calibrated by Beam Base Alignment.
BBA can eliminate (to a certain extent) the offset between the electric center and the magnetic center but can not eliminate the coupling. x=A0+A1Xraw+A2Yraw
Reviewers’ Comments from the CEPC_SPPC Pre-CDR Mini-Review (October 13-17, 2014, IHEP)

## Slide 45: 8-electrode BPM
e
p
Qp, xp, yp
Qe, xe, ye
8 Equations, 6 unknown
Tan α×L
=0.0165 ×740 =12.2 mm
Poor accuracy for off-axis beam
The Signal of the BPM with two beam
Kotaro Satoh,Detection of the position of two beams with a common BPM, PAC 1999

## Slide 46: 8-electrode BPM simulation
15
±3mm ×±15 mm
e+e- ：±12.2 mm
A
B
C
D
H
G
F
E
20
CST simulation confirmed the difficulty of off center beam measurement

## Slide 47: Study on BPM support
Material with low thermal expansion coefficient INVAR, carbon fiber, thermal expansion coefficient 1.2E-6/K, first-order characteristic frequency ＞ 50 Hz*.
* Wang anxin, Structural optimization design for BPM support of HEPS, High power laser andparticle beam, Vol. 33 No.4