## Slide 1: CEPC IP Beam position monitor
On behalf of beam instrumentation group
March 30- April 1, 2023
Hunan, Hengyang

## Slide 2
Contents
The function of IP BPM
1
The challenge for IP BPM
2
IP BPM preliminary design
3
Summary
4

## Slide 3: The principle of beam position monitor
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-

## Slide 4: 2 / 4 electrodes BPM
Principle of BPM: the closer the beam to the electrode, the stronger the signal, the position is proportional to the difference between electrodes
2 electrodes: x0= K× (R-L) /(L+R) K: the scale coefficient
Beam
x0
A
B
C
D
4 electrodes: x=Kh×U，y=Kv×V,
x0=y0=1
L
R

## Slide 5: IP BPM Vs other BPM
The IP Beam Position Monitor can also be used as a beam arrival time detector except position measurement, the beam-beam deflection technique is necessary for luminosity optimumization.
Phase advance
Phase lag
Spatially: The local bumps at the IP created by steering magnets can be used for beam–beam deflection scan. The parameters (position and angle) of electron and position can be set as the ideal state after through several iterations.
Temporally: The relative timing can be adjusted by changing the RF phase between the two rings, the collision point can be shifted.

## Slide 6: IP BPM: Temporally coincide
t3 – t1 = Δt = 2×L/c
L
L
t1
t2
t3
BPM location

## Slide 7: The challenge for IP BPM
1：~300 nm resolution and ~5μm accuracy
2：Limitation induced by material and installation space
3：Limitation induced by wakefield
4：Limitation induced by interfere for duel-beam
5：Other restrictions such as radiation resistance, beam is at nonlinear region

## Slide 8: Constraints on space and material
beam
air
vacuum
Seal contact face
Ceramic metallization
Safe distance

## Slide 9: Key components: feedthrough
Machine drawing
Inner structure results of x-ray tomography

## Slide 10: 4 Feedthrough + pipe = BPM pickup
Installation space requirement: 10 mm in longitudinal direction, 30 mm in radial direction
Welding between the feedthrough housing and the beam pipe, optional materials for housing: stainless steel, Kovar (nickel cobalt ferrous alloy), titanium, etc

## Slide 11: The wakefield of BPM
The signal of BPM electrode outputs
The longitudinal wakefield potential of BPM
The electric field in bpm during beam passby
A high-order mode captured by feedthrough

## Slide 12: Design of feedthrough
Reduce the volume of the gap and the sealing material
A lower dielectric sealing material, glass instead of ceramics
A asymmetric structure
Transmission structure optimization, a bigger coupling a smaller Qext
Special structure …
1
2
f ∝1/(εr)0.5，ceramics: 9.5 glass: 3~5.
Three main parameters for BPM
Radius and height of button rb，tb，gap gb，a set of typical parameter (rb=3mm, gb=0.25mm, tb=3 mm)，fH110 : 5.1 GHz, 7.6 GHz
2
Smaller the better (Mechanical processing)

## Slide 13: Design of feedthrough
3
3
offset
5
5
4
4

## Slide 14: Signal interference
Time interval between two beam pass through the same BPM ΔT : 2×L/c
CEPC IP BPM （s=640mm）ΔT: 4.26 ns, Signal width after the filter T1>Bunch gap T2>ΔT
4 electrodes
BII 8 electrodes BPM signal

## Slide 15: The challenge in a shared pipe
Operation mode of BPM readout electronics: CW mode Sampling rate ~100MHz.（different with scope GHz）
Bunch gap: 48 ns
Bunch gap: 24 ns
Gap: 12 ns
The filter will bring a long tail which make the detection become difficult

## Slide 16: The accurancy for off axis beam
＜200 μm
＜100 μm
＜60 μm
＜30 μm
The error distribution in a round pipe (x = A0+kU+A2V n=1)
n=1
＜100 μm
＜100 μm
＜50 μm
＜100 μm
n=3
n=5
n=7
IP BPM b=10 mm
BPM b=28 mm

## Slide 17: The error distribution
15
A
B
C
D
H
G
F
E
20
e+e-：±10.6 mm
Δt = 4.26 ns
±3mm ×±15 mm

## Slide 18: IP BPM distributions
4 electrodes: blue / red in separated pipe for position measurement
8 electrodes: black in shared pipe for position/arrive time measurement
s = 885 mm
s = 915 mm
s = 630 mm
Improve the position measure accurancy for a single beam but bring difficulty to separate the duel-beam signal
Relativity between the data and the beam at IP became less

## Slide 19: IP BPM
Preliminary schematic
Courtesy of Wang Haijing
A special pipe （thicker），shift 30 mm in Z for two beam.

## Slide 20: Mechanical parameters
 |  | Increase | Advantage | Disadvantage
rb | 3 mm | Larger area and Cb | Higher signal level | Lower resolution
tb | 2 mm | Larger volume | Lower wake impedance | Lower resolution
gb | 0.25 mm | More trapped modes | Higher resolution | Higher wake impedance
b | 10 (28) mm | Smaller signal | Lower wake impedance | Lower resolution
Solution for a thin pipe
Schematic for a track pipe

## Slide 21: Design parameters
 | Bunch charge | Bunch length | Current peak | Current average | 650MHz | Kloss
 | nC | mm | A | A | dBm | mV/pC
Higgs | 20.8 | 4.1 | 607 | 0.208 | -11.9 (-20.2) | 6.76 (0.16)
Z | 22.4 | 8.7 | 308 | 0.974 | 1.5 (-6.8) | 0.30 (0.15)
W | 21.6 | 4.9 | 527 | 0.216 | -11.5 (-19.8) | 3.4 (0.25)
ttbar | 32.0 | 2.9 | 1315 | 0.320 | -20.2 (-16.4) | 15.1 (0.017)

## Slide 22: Plan for next step
Signal high-speed switching processing
Algorithm improvement for 8 electrodes（with the help of calibration）
Development of glass sealing feedthrough
Prototype production including feedthrough and BPM
Wakefield Impedance test for a real BPM
….

## Slide 23: Summary
The performance of the machine depends on the performance of the measurement system.
The resolution and accuracy of BPM are constantly improving with the progress of mechanical processing and digital signal processing technology.
IP BPM will face a series of special issues, the spatial interference, background noise, algorithm optimization, wakefield heating, and welding process need to be solved and further research is needed.

## Slide 24
Thank you for your attention! 

## Slide 25
其它测量的分辨率精度、可重复性

## Slide 26: 高分辨率、高精度
 | BEPCII (800/200μm) | HEPS (10/1 μm) | CEPC
分辨率 | ＜ 10 μm （2s） | 0.1 μm | 1 μm (4 s)
精度 | 100 μm | /（10 μm） | 10 μm