## Slide 1: AHCAL MIP Calibration
Yukun Shi

## Slide 2: Motivation and strategy
Use MIP data and Simulation information to do energy calibration
Analyse HCAL and ECAL+HCAL muon data separately
  HCAL: muon beam position fixed
  ECAL+HCAL: muon beam position scanned
First do fit on chip level
  for channels with adequate counts, do fit on channel level
  for channels without enough statistics, use chip results as itself parameter

## Slide 3: Statistics overview
HCAL alone statistics
  There are channels with light leakage
  Most of the channel’s statistics are around

## Slide 4: Statistics overview
HCAL alone leaked channels
  No obvious time dependence
  Abandon channels with > counts
abandoned

## Slide 5: Statistics overview
Some abnormal NDL channels
Selection criteria:
  More than counts & Mean value < 250
  90 channels(0.4%) are abandoned
Abnormal NDL channel
Hit Map after selection

## Slide 6: Statistics overview
Data in the ECAL+HCAL run
  The light leakage has been fixed
  NDL layers still have abnormal channels
  The statistics is not enough for fittng at channel level

## Slide 7: MIP Selection
Some events have showers
Saturation has also been observed

## Slide 8: MIP Selection
Shower rejection
  Max layer hits < 4
  4292515 of 6453816(67%) passed
Fit in chip level is finished

## Slide 9: Scan of the
increased from 20 to 180 by 40
The peak value has an obvious increase

## Slide 10: Fit method test
increased from 20 to 180 by 40
decreased from 400 to 340 by 15
The fitted MPV is correct as if the fit is good

## Slide 11: Chip fitting results
3 chips in NDL layers fail the fitting
They will use layer

## Slide 12: Channel overview
Inadequate statistics usually will cause abnormal fitting result
Small MPV means the MIP peak is not well distinguished from pedestal which is not so reliable
Abnormal MPV are usually accompanied with other abnormal parameters
Inadequate statistics
Small MPV

## Slide 13: Channel overview
Large Gauss Sigma means the whole fitting is wrong even if other parameters seems normal
Abnormal landau sigmas gather up in NDL layers, mostly caused by inappropriate detector setup
Large gauss sigma
Abnormal landau width

## Slide 14: Channel overview

## Slide 15: Channel selection
Selection Criteria
  Entries>700
  < 20
  200<MPV<400 || ||<80
  20<Landau Width<100
  10<Gauss Sigma<150
Channels fail the selection will use chip MPV
Tag: 1 for normal; -1 for abnormal chips; -2 for inadequate statistics; -3 for abnormal fitting

## Slide 16: Summary
Event Selection has been studied to extract the light leakage events and shower events
The study of Landau gauss function and pre-fit on chip level guarantee the stability of the fitting
The MIP parameters for all channel is calibrated after the final selection on channels

## Slide 17
backup

## Slide 18: Energy Calibration
Calibration File V1
  Pedestal: cosmic_pedestal_beam.root
  GainRatio: Mu_Calib.root
  MIP: MIP_Xu.root
Planned V2
  Pedestal: DAC.root
  GainRatio: Pion_Calib.root
  MIP: Combined Xu,Hu,Shi result

## Slide 19: Different threshold
Muon beam
330 DAC
Peak 700
320 DAC
310 DAC

## Slide 21
83 84 87 36 37 38

## Slide 22: 漏光通道

## Slide 23
39 38 layer被误去除一部分

## Slide 24: MIP Selection
Some muon shower

## Slide 25: Statistics overview
Some abnormal NDL channels
Selection criteria:
  More than counts & Mean value < 150
  57 channels(0.4%) are abandoned
Abnormal NDL channel