## Slide 1: Design, Construction and Operation of the BESIII Detector
Yifang Wang
Institute of High Energy Physics

## Slide 2: Origin of the BESIII
From BES, BESII and Tau-Charm factory
Experience from Babar, BELLE, CLEO/CLEO_c, KLOE etc.
Efforts by several generations of physicists
Goal: Much smaller systematic uncertainties(full coverage, good resolution, etc) and much higher data rate （~3 kHz @ J/y)

## Slide 3: Design Evolution of BESIII
For cost considerations:
Original baseline design: A scintillating fiber (KLOE type) EM calorimeter with existing BESII magnet 
Superconducting magnet was an option if there are international contributions
A new possibility: L3 BGO crystals refurbished as the EM calorimeter 
CsI(Tl) crystals 24 cm long  28 cm long
Once we experienced a luxury life, we cannot go back: CsI(Tl) (28 cm) + Superconducting magnet by cost saving + budget increase

## Slide 4: Main Mitigation Measures
The process was non-linear, with many iterations. A simplified summary:
Reduced the CsI(Tl) cost from 2.5$/cc to <1.8$/cc, at the support of SIC, thanks to Prof. Dongsheng Yan(严东生) and Prof. Renyuan Zhu ( ~5M$)
A hard decision to build the superconducting solenoid magnet by ourselves (cost  3, ~6M$)
Removed the scintillating fiber vertex detector (originally for trigger)
Reduced the budget for trigger and electronics (~1M$)
Budget increase of ~ 4M$
Design was started in ~1999 and finalized at ~2004

## Slide 5: Major Design Issues
Overall mechanics design by Prof. Xuan Liu (and Shaowang Xu)
Superconducting magnet design helped by KEK
Innovative design of the supporting structure for the CsI(Tl) Crystal calorimeter by Prof. Lingshu Wang: using screws instead of honeycomb structure to hold crystals: much cheaper(>1M$) and much less dead materials between crystals
MDC: no outer wall for inner chamber and no inner wall for outer chamber to reduce materials: Two chambers are glued together at the end plates

## Slide 6: Construction: Main Drift Chamber
Led by Prof. Yuanbo Chen, electronics by Prof. Huayi Sheng
Main R&D:
  Materials: wire(tension, creeping, …), glue(aging, radiation hardness), feedthrough(precision, HV resistance), …
  Prototyping & test beam
Manufacturing:
  Precision of end-plates: failed twice and wondering around for a year
  Feedthrough: precision, contract dispute, …
  Wiring, carbon fiber cylinders, precision of assembly, gas tightness, …

## Slide 7: Construction: CsI(Tl) Crystal Calorimeter
Led by Prof. Junguang Lu, electronics by Prof. Peiliang Wang
R&D: light yield, uniformity, wrapping, PD, test beam,…
CsI(Tl) crystals produced from SIC, Hamamatsu and Saint-Gobain in order to control the risk, cost and quality
Low noise electronics: 1000 e equivalent ~ 220 keV

## Slide 8: Superconducting Magnet
Led by Prof. ZiAn Zhu, in collaboration with WANG NMR INC.
Issues: structure, cable winding, assembly, valve box…
Huge risk: not possible to test before the final curing
  too late if we know it is not good
After 3 times of quenching, the magnet reached 1T
Field mapping: Dr < 0.5 mm, DB < 0.1%
1T magnet @3364A

## Slide 9: Construction: Time-Of-Flight
Led by Prof. Yuekun Heng, electronics by Prof. Qi An(USTC)
Careful choices of the Plastic scintillator and PMTs
R&D: radiation hardness, wrapping, optical coupling, effects of magnetic field, test beam, …

## Slide 10: Construction: RPC muon chamber
Led by Prof. Jiawen Zhang, electronics by Prof. Yongzhao Zhou（USTC）
Invented a new type RPC made of Bakelite with a melamine surface. No linseed oil to avoid aging problem
Noise less than 0.1 Hz/cm2
RPC surface seen by the atomic force microscope
IHEP
Babar w/o oil
Babar w/ oil
Belle glass
C. Lu, Princeton University

## Slide 11
Electronics, Slow Control and Grounding
Drift chamber：6500ch，st ~ 500ps, sq ~ 5 fc, 10bit ADC
calorimeter: 6300ch, sq ~ 0.5fc, 310bit ADC, noise < 1000ENC
TOF：500ch，st ~ 20 ps, 10bit ADC
RPC：10000ch, bit map
No major failures up to now, nor any need to major upgrade
Cleaning grounding was carefully designed to be isolated from dirty ground but:
  Ground of BESIII and BEPCII were connected by the cryogenic pipe for BESIII SC magnet
  RF signals may still couple to MDC from time to time
IHEP
USTC

## Slide 12
Trigger System
RD
TCL
BUS
RD
TCL
BUS
RD
TCL
BUS
RD
TCL
BUS
RD
TCL
BUS
Fan out
Main trigger controller
Trigger condition processor
Trigger table
TDC
MDC
MDC
EMC
TOF
CLK
NIM crate
Scaler
9UVME
6UVME
From front electronics
logic
To fiber
backplane
Self-defined
bus
Trigger condition
Scaler signal
CLK input
Crate CLK
Crate connection
Trigger controller
fiber
A compact system with：
4× 9UVME
1×6UVME
1×NIM
Based on RocketIO, FPGA and optical fiber technologies (late used for BELLE II and Panda)
scaler

## Slide 13
Beam related Background
Backgrounds:
SR photons
Beam-gas
Bremsstrahlung
Coulomb scattering
Mitigations:
Beam optics design
Gold plating of Be pipe
Masks
Movable masks
Results:
No issues after careful design and manufacturing

## Slide 14
Be Beam-pipe
Led by Prof. Quan Ji
Original plan was to buy from Brush Wellman, not possible later on
Design:
  30 cm long Be pipe with 15 mm Gold plating to stop SR photons
  2 layers of Be pipe with oil cooling channel in between, coated with 20 mm BR127
  Total heat to be removed: HOM ~200W, SR < 24 W
R&D: structure, Be pipe manufacturing, cooling gap and material, gold plating, Be-Al welding, …

## Slide 15
Installation
MDC+TOF clearance < 10 mm
EMC clearance < 15 mm

## Slide 16
BESIII software and data flow
Slow control
database
Slow control farm
Beam/envion. para.,
Luminosity, dose,
HV, Tem.…
Raw Data
stream:
Online farm
Online
database
detector
DAQ;
Event filter
Data Quality monitoring;
Event display
Offline
database
Offline farm
MC generator
MC simulation
Reconstruction
Calibration
REC/DST
Data stream
RAW/REC data:
Bhabha,
Dimuon
Hadron,
Beam gate …
Physics
tools
Physics
analysis

## Slide 17
DAQ and Offline
Led by Prof. Kejun Zhu
specification：~ 50Mb/s, 4000 Hz, 10 B-factory, 1000  BESII
Online monitoring，display and control
Transient Event Store
Generator
Simulation
Digitization
Gaudi
Event Display
Simulation
Reconstruction
Physics Analysis
Calibration
Cal-MDC
EVERTEX
EMC(B/E)
V0FIND
Exp. data
MC data
Detector-Geo
Cal-Data
Control-Data
DST
Made
Cal-TOF
Fast-track
Event-T0
MUID
TOF(B/E)
MDC-track
REC
Data
DST
Data
Pre-tracking
Post-tracking
Phy-tracking
calibUtil Services
Register, write
Search
Read
Gaudi
Client
Calibrator
Gaudi
services
Metadata (persistent)
MySQL rdbms
FACS
Calibration &
Database
Simulation
Reconstruction
Led by Prof. Weidong Li
CMT based BESIII software development and management
GAUDI based software framework
GEANT4 based Monte Carlo simulation
MYSQL based calibration and database

## Slide 18
Quality Control
Rules and regulations
Documentation
Education
Process control
Regular inspection
Review

## Slide 19
Cosmic-ray Test
Debug hardware problems
System tests
Detector Calibration
Results
  MDC single wire resolution:
    < design spec. 110%
  EMC energy resolution
    < design spec. 140%
  TOF time resolution
    < design spec. 130%
  MUC efficiency
    < design spec. 100%
We are ready for the first physics results

## Slide 20
Commissioning Milestones
Nov. 18, 2007: first e+e- collision
Mar. 2008: Collision 500 mA × 500 mA, Lumi.: 1 ×1032 cm-2s-1
April 30,2008: Move BESIII to IP
July 19, 2008: First event in BESIII
BESIII at the interaction point
Precision < 1 mm

## Slide 21: Performance of the Detector
Exps. | MDC Wire reso. | MDC dE/dx Reso. | EMC Energy reso.
CLEO | 110 m | 5% | 2.2-2.4 %
Babar | 125 m | 7% | 2.67 %
Belle | 130 m | 5.6% | 2.2 %
BESIII | 115 m | <5% (Bhabha) | 2.3%
Exps. | TOF time reso.
CDFII | 100 ps
Belle | 90 ps
BESIII | 68 ps (Barrel) 98 ps (Endcap) 60 ps(Endcap, MRPC)
MUC: Eff. ~ 96%
Noise: < 0.04 Hz/cm2(Barrel), < 0.1 Hz/cm2(Endcap)

## Slide 22: Stable Operation for 15 years
MDC :
  No broken wire. 33 dead channels in total, 30 channels can be repaired due to preamp issues
  Hit efficiency: 90%~97%, except for the first 3 layers
EMC:
  No dead crystal module. 27 unglued modules，light yield of 5 modules dropped to about 60-70%
  Efficiency : Almost 100%, except for the 5 unglued modules
BTOF:
  No dead channel
  Efficiency : 95%~97% for single end
ETOF:
  6 dead channels can be repaired due to preamp issues
  Efficiency : 98.5%, except for the dead channels
MUC:
  1 module dead (east endcap layer0)
  Efficiency : >90%
MDC hit efficiency vs layer
BTOF efficiency
No major failures

## Slide 23: Integrated luminosity： 46.58 fb-1

## Slide 24
Collaboration
Drift chamber and its electronics (IHEP, Sichuan, Tsinghua)
CsI(Tl) calorimeter and its electronics(IHEP, Tsinghua)
TOF（IHEP, USTC，Tokyo，Hawaii）
TOF electronics （USTC）
RPC（IHEP, Uni. of Washington）
RPC electronics（USTC）
Trigger（IHEP, USTC, Giessen）
DAQ & online software (IHEP, Tsinghua)
Offline software （IHEP, Peking，Shangdong，Nanjing，JINR）
Superconducting magnet （IHEP）
Mechanics(IHEP)
Technical support ( IHEP, Tsinghua,BINP）
Japan (1)
Tokyo University
USA (2)
University of Hawaii
University of Washington
China (22)
IHEP, CCAST, GUCAS，
Univ. of Sci. and Tech. of China
Shandong Univ., Zhejiang Univ.
Huazhong Normal Univ., Wuhan Univ.
Zhengzhou Univ., Henan Normal Univ.
Peking Univ., Tsinghua Univ. ,
Zhongshan Univ.,Nankai Univ.
Shanxi Univ., Sichuan Univ
Hunan Univ., Liaoning Univ.
Nanjing Univ., Nanjing Normal Univ.
Guangxi Normal Univ., Guangxi Univ.
Europe (4)
GSI, Germany
University of Bochum, Germany
University of Giessen, Germany
JINR, Dubna, Russia
New BESIII collaboration was established in Jan. 10, 2006
~300 members
From 29 institutions in 5 countries

## Slide 25
China (55/367)
Institute of High Energy Physics (146), other units(221): Beijing Institute of Petro-chemical Technology, Beihang University,
China Center of Advanced Science and Technology, Fudan University,
Guangxi Normal University, Guangxi University,
Hangzhou Normal University, Henan Normal University,
Henan University of Science and Technology,
Huazhong Normal University, Huangshan College, Hunan University,
Hunan Normal University, Henan University of Technology
Institute of modern physics, Jilin University, Lanzhou University, Liaoning Normal University, Liaoning University, Nanjing Normal University, Nanjing University, Nankai University, North China Electric Power University,
Peking University, Qufu normal university, Shanxi University,
Shanxi Normal University, Sichuan University, Shandong Normal University,
Shandong University, Shanghai Jiaotong Univeristy, Soochow University,
South China Normal University, Southeast University, Sun Yat-sen University,
Tsinghua University, University of Chinese Academy of Sciences, University of Jinan, University of Science and Technology of China,
University of Science and Technology Liaoning,
University of South China, Wuhan University, Xinyang Normal University,
Zhejiang University, Zhengzhou University,YunNan University , China University of Geosciences , Yantai University,
Europe (17/115)
Germany (6): Bochum University,
GSI Darmstadt, Helmholtz Institute Mainz, Johannes Gutenberg University of Mainz, Universitaet Giessen,University of Münster Italy (4): Ferrara University, LNF/Perugia, University of Torino Netherlands (1):KVI/University of Groningen
Russia (2): Budker Institute of Nuclear Physics, Dubna JINR
Sweden (1):Uppsala University
Turkey (1):Turkish Accelerator Center Particle Factory Group UK (2): University of Manchester, University of Oxford
Poland (1)National Centre for Nuclear Research
USA(4/8) Carnegie Mellon University Indiana University University of Hawaii University of Minnesota
~600 members
From 83 institutions in 16 countries
Asia (6/10)
Pakistan (2): COMSATS Institute of Information Technology
University of the Punjab, University of Lahore
Mongolia (1): Institute of Physics and Technology
Korea (1): Chung-Ang University
India (1): Indian Institute of Technology madras
Thailand (1): Suranaree University of Technology
South America (1/1) Chile: University of Tarapaca
Current Collaboration

## Slide 26
Preparation for Physics
conferences：
  Charm series
  China-US collaboration workshops
  Summer schools
International Workshop on
tau-Charm Physics
Beijing, June 5-7, 2006
hep-ex/0809.1869
IJMP A V24,No 1(2009) supp
Charm2023 Siegen (Germany 2023)
Charm2020 Mexico City (Mexico 2020)
Charm2018 Novosibirsk (Russia, 2018)
Charm2016 Bologna (Italy, 2016)
Charm2015 Detroit (USA, 2015)
Charm2013 Manchester (UK, 2013)
Charm2012 Honolulu (USA, 2012)
Charm2010 Beijing (China, 2010)
Charm2009 Leimen (Germany, 2009)
Charm2007 Ithaca (USA, 2007)
Charm2006 Beijing (China, 2006)

## Slide 27: Thanks to IHEP/BEPCII management, in particular, Prof. Weiguo Li as the deputy manager for BESIII
And
Chief engineers Prof. Xuan Liu (and Prof. Shaowan Xu)
Chief technical support staff: Prof. Jingzhi Bai
All L2 & L3 sub-system managers
Run coordinators
And all collaborators