## Slide 1: The radon and radium measurement systems for JUNO
Cong Guo IHEP
On behalf of the JUNO-VETO group
2023-11-07

## Slide 2: Outline
An introduction to JUNO
Radioactive background requirements on ultra-pure water
Radon measurement and radon removal
Radium measurement and radium removal
The status of the 100 t/h ultrapure water system for JUNO
Summary

## Slide 3
JUNO: A multipurpose neutrino experiment
Phyiscs:
Determine neutrino mass ordering
Precise measurements of three neutrino oscillation parameters
Supernova neutrinos
Solar neutrinos
Atmospheric neutrinos
Geo-neutrinos
Exotic searches, proton decay

## Slide 4
JUNO: A low-background experiment
Research | Expected Signal | Energy range
Reactor neutrino | 60 IBD/day | 0~12MeV
Solar neutrino | Hundreds per year for 8B | 0~16MeV
Geoneutrino | ~400 per year | 0~3MeV
Atmospheric neutrino | Hundreds per year | 0.1~100GeV
DSNB (w/s PSD) | 2-4 IBD/year | 10-40MeV
Supernova burst | 5000 IBDs at 10kpc 2300 elastic scattering | 0-80MeV
The expected signal event rate is low and the energy is small;
The background of the detector should be well-controlled;
The detector is located 700 m underground and submerged in a Water Cherenkov Detector (WCD).

## Slide 5
JUNO signal and background
Reactor neutrino detection (IBD):
Prompt signal: 0.7 MeV< Ep < 12 MeV;
Delayed signal: 1.9 MeV < Ed < 2.5 MeV;
Natural radioactive elements (214Bi from 238U-226Ra-222Rn decay chain, 208Tl from 232Th decay chain ) and cosmic muons are the main background sources;
Solar neutrino detection (ν-e elastic scattering):
Signal as a single event;
Events in 2-3 MeV dominated by 214Bi (238U-226Ra-222Rn decay chain) decays and 212Bi (232Th decay chain) decays;
Events in 3-5 MeV dominated by 208Tl (232Th decay chain) decays;
Estimated base on 238U/232Th in LS is 10-17 g/g;
Internal radioactivity background after cuts for solar neutrino detection
IBD

## Slide 6
Background requirements on UPW
LS filling
Ultra Pure Water (UPW) in the WCD:
  Gamma rays from the radioactive elements in water will contribute to the accidental background;
UPW for LS filling:
  LS contacts UPW directly;
  Water may pollute LS;
 | Veto | Filling
Flow rate | 100 t/h | 50 t/h
238U/232Th | 10-14 g/g | 10-16 g/g
222Rn concentration | 10 mBq/m3 | 5 mBq/m3
226Ra concentration | 10 mBq/m3 | 50 μBq/m3
Rn detector sensitivity | 1 mBq/m3 | 1 mBq/m3
Ra detector sensitivity | 1 mBq/m3 | 10 μBq/m3

## Slide 7
Rn concentration in water measurement
Atomizer
Schematic diagram of the measurement system
Radon in water can not be measured directly;
An atomizer is used to transfer radon from water into gas,
  CRn in water/CRn in gas =

## Slide 8
Rn concentration in water measurement
The radon concentration is determined by detecting the alphas from 218Po and 214Po decay using a Si-PIN detector;
The detecting efficiency is calibrated with a gas flow radon source;

## Slide 9
Rn concentration in water measurement
A semi-automatic radon concentration measurement system in water has been developed;
Electromagnetic valves are used in the system;
Software has been developed for remote operating;
Based on the background and the calibration results, the sensitivity of the system is ~1mBq/m3;
Atomizer
Rn detector
Radon concentration in water measurement system

## Slide 10
Rn removal from water
Schematic diagram of degassing membrane
Working principle of Micro-bubble device
The main device for removing radon from water is the degassing membrane, which is made up of microporous hollow fibers;
The radon removal efficiency decreases as the gas in the water decreases;
Micro-bubble device is used to load gas into water, so as to increase the radon removal efficiency;
According to the prototype measurement results, the radon concentration in water has been reduced to 3.6 ± 2.3 mBq/m3 from ~1Bq/m3;

## Slide 11
Ra concentration in water measurement
226Ra, the progenitor of 222Rn, has a half-life of 1600y, 10 μBq/m3 ≈ 5* 10-22g/g;
MnO2 has strong adsorbability for radium;
Mn-fiber is used to extract radium from water, Mn-fiber is polyurethane fiber with MnO2 attached;
The radium extraction efficiency is calibrated with 226Ra solution;
Extraction column
Mn-fiber
226Ra solution

## Slide 12
Emanation chamber
Ra concentration in water measurement
226Ra is detected by its gaseous daughter 222Rn;
After several days of sealing, the 222Rn is transferred to the radon detector to measure its activity;
The activity of 226Ra to 222Rn flows :
  ARa = ARn/(1-e-λt)
  t is the sealing duration, if t = 3.8 days, ARn= 0.5*ARa;
According to the background and calibration measurement results, the sensitivity of this system is ~10 μBq/m3;

## Slide 13
Ra removal from water
Radium exists in ultrapure water primarily in the ionic form;
The UPW production process is essentially a deionization process;
Multimedia filter, Revise Osmosis (RO), Electro-De-Ionization (EDI), and resin can be used to remove radium from water;
The prototype measurement results showed that the radium concentration can be reduced to 53 ± 40 μBq/m3 from ~500mBq/m3;
The flow chart of a UPWS

## Slide 14
The UPWS of JUNO

## Slide 15
The UPWS of JUNO
Water room on the ground
Water system on the ground
Water system underground
The surface building was completed in 2022;
The installation of the water system is basically finished, only some pipes need to be connected;
The commissioning of the system is expected to start in Dec. 2023.

## Slide 16
Summary
JUNO is a multi-purpose neutrino experiment
JUNO is also a low-background experiment under construction at 700m underground and submerged in a WCD
In addition to being used in the WCD, UPW will also be used for LS filling
An UPWS is developing to produce very low radioactive UPW (222Rn < 5mBq/m3, 226Ra <50 μBq/m3 ) with a flow rate of 100t/h;
Two setups for measuring 222Rn (226Ra) concentration in water with a sensitivity of ~1 mBq/m3 (~10 µBq/m3) have been developed
Based on the prototype measurement results, the radon concentration in water can be reduced to 3.6 ± 2.3 mBq/m3 and the radium concentration in water can be reduced to 53 ± 40 μBq/m3
The installation of the UPWS is ongoing and the commissioning is expected to start in Dec. 2023