## Slide 1: Construction and Test of Prototype of Advanced Multi-neutron Detection Array
Presenter: Jiawei BIAN
Supervisor: Zaihong YANG
State Key Laboratory of Nuclear Physics and Technology, Peking University
1

## Slide 2: Outline
Background
Construction and Cosmic Ray Test
  Setup
  Data Analysis
Collimated Proton Beam Test
Summary
2

## Slide 3: Background
3
neutron halo
neutron correlation and neutron cluster
cluster
unstable nuclei
stable nuclei
shell model
neutron star
exotic structure

## Slide 4: Background
Preliminary design of AMDA
(Advanced Multi-neutron Detection Array)
50 units each plane, 30 planes
Adjacent planes arranged vertically
4
Notes: Granularity
To verify some technique index
Describe the signal readout
Multi-neutron efficiency and capability

## Slide 5: Outline
Background
Construction and Cosmic Ray Test
  Setup
  Data Analysis
Collimated Proton Beam Test
Summary
5
Notes: PMT->SiPM a more complete construction and test

## Slide 6: Setup
Prototype array
DAQ (CAEN)
  DT5730
  PC(CoMPASS)
6

## Slide 7: Setup: Detection Unit
Saint-Gobain BC408 plastic scintillator
(with ESR)
SiPM: SensL-C60035
SensL-C60035
Senser size | Microcell size | Parameter | Overvoltage | Typ. | Unit
6mm | 35μ | Breakdown Voltage |  | 24.2-24.7 | V
- | - | Over voltage |  | 1.0-5.0 | V
- | - | Rise Time |  | 1.0 | ns
- | - | Gain |  |  | 
Senser size | Microcell size | Parameter | Overvoltage | Typ. | Unit
6mm | 35μ | Breakdown Voltage |  | 24.2-24.7 | V
- | - | Over voltage |  | 1.0-5.0 | V
- | - | Rise Time |  | 1.0 | ns
- | - | Gain |  |  | 
7
Notes: After attach…
Wrapped them entirely with shading tape

## Slide 8: Setup: Logic Settings
Self-trigger for each unit
4-units
coincidence
8
offline
Notes: make coincidence of A and B

## Slide 9: Outline
Background
Construction and Cosmic Ray Test
  Setup
  Data Analysis
Collimated Proton Beam Test
Summary
9

## Slide 10: Raw Data Check
Time difference and Energy spectrum for 1 unit:
10
Notes: Independent with x
Decay with x

## Slide 11: Calibration: Signal velocity
Differential Time difference gives out width .
Signal velocity given by and Length :
11
cosmic ray
Notes: Edge->fastest position ->peak in differential spectrum
Correspond to

## Slide 12: Calibration: Tracking
0: mid-point of bar; others: caused by module offset and so on
12
Notes: Consistent with the

## Slide 13: Position & Time Resolution
With the incident angle by tracking,
we can also select events with nearly vertical incidence:
x(cm)
13

## Slide 14: Attenuation Length
Energy of light signal decays with distance in plastic scintillator,
Attenuation length was derived by:
Cosmic rays
14

## Slide 15: Attenuation Length
has satisfied the need for experiment.
fit result of #bar2
15
Notes: Technique is still not mature

## Slide 16: Outline
Background
Construction and Cosmic Ray Test
  Setup
  Data Analysis
Collimated Proton Beam Test
Summary
16
Notes: In addition to …

## Slide 17: Collimated Proton Beam Test: Setup
proton beam
17
@CIAE, Beijing
steel tube
plastic scintillator
Beam energy: 17 MeV, 4 MeV
12C target
Notes: Get collimated proton
2 different energy
In scattering chamber

## Slide 18: Collimated Proton Beam Test: Result
20 cm
Signal velocity could be given with time difference between adjacent peaks and
18
Notes: 加标题下划线

## Slide 19: Collimated Proton Beam Test: Result
Beam energy(MeV) | Signal velocity(cm/ns) | Time resolution(ps) | Position resolution(cm)
17 | 13.31 | 152.55 | 1.438
4 |  | 149.38 | 1.406
19
Notes: >1k comment

## Slide 20: Summary
Prototype array constructed and tested
Time and position resolution, attenuation length have reached relatively high level.
Future plan：
Changing from CAEN DAQ to TOFPET-based electronics
Optimizing work conditions of SiPM
Developing machine learning algorithm for multi-neutron identification
20
(Zeyu DU et al. multi-neutron detection based on machine learning)
Thank you!

## Slide 21: backup

## Slide 22: backup
SiPM (Silicon Photon Multiplier)：