## Slide 1: Development of Polarized 3He at CSNS
2025.09.23 Qingdao
Junpei (Jupiter) Zhang
Institute of High Energy Physics (IHEP)
China Spallation Neutron Source (CSNS)
Guangdong Provincial Key Laboratory of Extreme Conditions
Guangdong Provincial Key Laboratory of Extreme Conditions

## Slide 2
Outline
I. Polarized Neutrons and Hyperpolarized ³He
II. Development and Applications of ³He-NSF
III. Conclusion and Future Plans

## Slide 3
Outline
I. Polarized Neutrons and Hyperpolarized ³He
II. Development and Applications of ³He-NSF
III. Conclusion and Future Plans

## Slide 4
CSNS-Dongguan
1.1 China Spallation Neutron Source
Introduction of China Spallation Neutron Source (CSNS)
Bird View of the Target Station
Bird View of the CSNS

## Slide 5
Neutron Instruments at CSNS
More than half of the neutron spectrometers (blue) at the CSNS have the demand or capability to equip with the function of neutron polarizing.
1.2 Polarized Neutron Demand at CSNS

## Slide 6
1.3 Polarizing Neutron with Polarized 3He
How Neutron Polarized by Polarized 3He
Including system:
Filling Station
Heating/Discharge
Filled Cells: Rb/K/3He
High Power Pumping Laser：795/1083 nm
Uniform Magnetic field
NMR/Compressor
……
Neutron Spin Filter

## Slide 7
Outline
I. Polarized Neutrons and Hyperpolarized ³He
II. Development and Applications of ³He-NSF
III. Conclusion and Future Plans

## Slide 8
3He Polarizing：SEOP & MEOP
SEOP
Advantages:
Integrable
Relatively Low Cost
Highly Experienced
Disadvantages:
Alkali Metals Required
High-Temperature: ~200 ℃
Low Production Rate: ~0.1 bar ·L /h
MEOP ：
Advantages:
Alkali Metals Free
Room Temperature：20 ℃
High Production Rate: ＞1 bar·L/h
Disadvantages:
Large Scare
Further Compress
Expensive
SEOP (Bouchiat et al., Phys. Rev. Lett. 5, 373, (1960) )
MEOP (Colegrove et al., Phys. Rev. 132, 2561, (1963) )
3He↑*
3He*
3He↑
3He
Rb
3He↑
Rb↑
3He
pHe ≈ 1-10 bar
pHe ≈ 1 mbar
Rb↑ + 3He = Rb + 3He↑
3He + 3He↑* = 3He↑ + 3He*
2.1 Introduction for Polarizing 3He
RF
23P0
23S1
11S0
F=1/2
F=3/2
F=1/2
F=3/2
F=5/2
F=1/2
F=3/2
F=I=1/2
1083 nm
C8
C9
Collisional Mixing
ms = -1/2
ms = 1/2
794.7 nm σ+
Quenching
N2
52S1/2
52P1/2

## Slide 9
2.2 3He Neutron Spin Filter (NSF)
Provide Polarized 3He-NSF for Neutron Spectrometers
Polarized 3He-NSF for CSNS
3He
Polarizing
SEOP
MEOP
In-situ
Off-situ
Off-situ

## Slide 10
High Power Pumping Laser
Pumping Station
Uniform Heating System
Uniform Magnetic Field
NMR System
3He NSF Cell
2.3 SEOP 3He-NSF
1. 3He-NSF R&D at CSNS

## Slide 11
2. 3He Cell Fabrication
≥180°
2. 3He Cell Blowing
4. Annealing
Wide-Angle
Cylindrical
1. Expansion
3. Stress Check
5. Stress Verification
6. Pressure Test
2.3 SEOP 3He-NSF

## Slide 12
3. Filling Station for 3He-NSF
3He Filling Station (2022)
GE180 Cells Production: ≥10-20/Year
Tip-off of 3He-NSF
P*L (barn*cm)
Neutron Wavelength (Å)
2.3 SEOP 3He-NSF
Cylindrical
Wide-Angle

## Slide 13
4. Pumping Station R&D
2.3 SEOP 3He-NSF
2019
2020

## Slide 14
Contents | Off-situ1 | Off-situ2
Power | ~ 160 W | ~ 130 W
Laser Type | Commercial | Home Build
Pumping Rate (τ) | 3-6 hours | 4-6 hours
Oven Size | ~ 8000 cm3 | ≥ 8000 cm3
Magnetic Filed | ≥ 12 G | ≥ 12 G
Gradient | ~ 4.2 ×10 -4 | ~ 4.2 ×10 -4
Hating Power | 200 ℃ 50 min | 200 ℃ 30 min
NMR | FID、AFP、EPR | 
Pumping Time：≤12 hours，Polarization：≥70%
4. Pumping Station R&D
2.3 SEOP 3He-NSF

## Slide 15
Polarized Neutron Experiment
Transportation
Pumping
Pumping time: 10-12 hours
Transport: ~30 minutes
Decay Constant: ≥100 hours
Continuing Decay→Short Experiment
Calibrating Required
Chinese Physics Letters, 38(9), 092801(2021).
77.4%
Decay Constant ≥202.9 hours
3He Polarization ≥ 77% and Decay constant ≥ 200 hours are achieved since 2020.
3He NSF
Neutron Beam
2.3 SEOP 3He-NSF
5. Off-situ 3He-NSF R&D

## Slide 16
pl ≈2.1
≈ 40 %
p ≈0.4~0.5 bar
6. Off-situ 3He-NSF-Wide Angle
2.3 SEOP 3He-NSF
Length ≈ 16 cm
Established the capability for wide-angle cell production at CSNS.

## Slide 17
Key Performance of the Current Off-situ Pumping System
System Type | System Length | Cell Length | Cell Pressure | 3He Polarization | Lifetime | Neutron Polarization （2 Å） | Neutron Transmission （2 Å）
Off-situ1 | ＜ 160 cm | ≥ 8 cm | ≥ 0.8 bar | 40%-77% | ≥ 100 h | ≥ 80% | ≥ 30%
Off-situ2 | ＜ 240 cm | ≥ 8 cm | ≥ 0.6 bar | 50%-75% | ≥ 100 h | ≥ 80% | ≥ 30%
Off-situ3 | ＜ 150 cm | ≥ 8 cm | ≥ 0.3 bar | 30%-75% | ≥ 100 h | -- | --
Chin. Phys. Lett., 38 (2021), DOI: 10.1088/0256-307x/38/9/092801
Off-situ 1
Off-situ 2
Off-situ 3
2.3 SEOP 3He-NSF
7. Off-situ 3He Pumping Stations

## Slide 18
Stable Polarization
Integrated、Mobile
Easy Data Analysis
In-situ 3He-NSF
2.3 SEOP 3He-NSF
8. In-situ 3He-NSF R&D

## Slide 19
In-situ SEOP NSF Provides Long-Time Stable Polarized Neutron
9. In-situ 3He-NSFs
China. Phys. Lett., 42, 022901 (2025).
System Type | System Length | Cell Length | Cell Pressure | 3He Polarization | Neutron Polarization （2 Å） | Neutron Transmission （2 Å）
In-situ1 | ＜ 70 cm | ≥ 10 cm | ≥ 1.5 bar | ≥ 74% | ≥ 90% | ≥ 20%
In-situ2 | ＜ 55 cm | ≥ 8 cm | ≥ 2.5 bar | ≥ 75% | ≥ 95% | ≥ 15%
In-situ3 | ＜ 60 cm | ≥ 8 cm | ~ 0.9 bar | ≥ 65% | ~ 80% | ≥ 20%
In-situ 1
In-situ 2
In-situ 3
In-situ 1+In-situ 2
Sci. China Phys. Mech. Astron., 65, 241011 (2022).
2.3 SEOP 3He-NSF

## Slide 20
1. MEOP System
2.4 MEOP 3He Pumping Station
3He Polarimetry
Optical Pumping
RF Discharge
Magnet
B0
OPC
3He Bottle
Purifier
Capillary
Optical Pumping Cells
Yb Fiber Laser
OPC
Buffer 2.5 L
5.2 L Compressor
Hydraulic Piston
Polarized 3He Cells
Optical Polarimeter
Mirrors
Discharge
1083 nm
Yb Fiber Laser
Holding field for 3He
Generate metastable 3He*
1083 nm circular polarized narrow line laser for 3He pumping
Calibrate the polarization of 3He
Compressor
Compress polarized 3He from mbar to ≥1 bar
Cell
Filling with 3He

## Slide 21
2. MEOP R&D
2.4 MEOP 3He Pumping Station
Cell
Design
Cleaning & Filling
Sealed Cell
Tip off

## Slide 22
2. MEOP R&D
2.4 MEOP 3He Pumping Station
Magnetic Field
Magnetic Field Optimization
Magnet System
Distribution of Magnetic Field & Gradient

## Slide 23
2. MEOP R&D
2.4 MEOP 3He Pumping Station
Optical System
Magnet
QWP
Pump Laser
Mirror
Probe Laser
Beam Shaping
QWP
Balanced Photodiode
RF Discharge
HWP
PBS
C9: 1083.327 nm
C8: 1083.356 nm
Light Trap
3He Cell
PBS
Balanced Photodiode
Mirror
Mirror
PBS
PBS

## Slide 24
3. Current Status
RF Electrode & Probe Laser
Ionization & Pumping
2.4 MEOP 3He Pumping Station
Current Status of MEOP

## Slide 25
2.4 MEOP 3He Pumping Station
Pumping Parameters:
3. Current Status
Pressure: ~1 mbar
Pumping Results:
Pumping Rate: 5 s/15 s
Decay Time: 195 s
Pumping Power: 4 W
Cell Diameter: 5 cm
3He Polarization: ~65%
Cell Length: 10 cm

## Slide 26
4. Next step Horizontal or Vertical
2.4 MEOP 3He Pumping Station
Horizontal （ILL Tyrex）
Vertical （ANSTO Flynn）

## Slide 27
In-situ NSF
Calibration of Neutron Spin Control Devices
CARR
MR
Back-n
NOPTREX
P-SANS
Neutron Beam
Neutron Beam
Sample
In-situ 3He NSF
Detector
3He Supports
SM
Neutron Beam
In-situ
3He NSF
Sample
Neutron Beam
Neutron Beam
Supports
In-situ
3He-NSF
Neutron Beam
Detector
Sample
2.5 3He-NSF Commissioning
Inelastic Spectrometer

## Slide 28
Multiple Calibration：
Supper Mirror
Mezei Flipper
RF Spin Flipper
Supper conducting
Flipper
Zero Field Chamber
··· ···
1. Calibrating R&D Devices
Supper Mirror
RF Spin Flipper
Off-situ 3He-NSF
In-situ
3He-NSF
Off-situ
3He-NSF
Supper conducting
Flipper
Supper conducting
Flipper
Supper conducting
Flipper
Off-situ 3He-NSF
Zero Field Chamber
2.5 3He-NSF Commissioning

## Slide 29
Polarized 3He：In-situ2
Polarization：≥68%
Pressure：2.53 bar
Beam diameter: Φ = 50 mm
2.5 3He-NSF Commissioning
2. 3He-NSF for NOPTREX @ Back-n (1st)
First use of a polarized 3He neutron spin filter on the Back-n White Neutron Source of CSNS

## Slide 30
2. 3He-NSF for NOPTREX @ Back-n (1st)
3He Polarization＞68%, Statablity~100 hours.
First use of a polarized 3He neutron spin filter on the Back-n White Neutron Source of CSNS, NIMA, 1072 (2025) 170184.
Polarized 3He
Un-Polarized 3He
Neutron Energy（eV）
Neutron Counts
3He Polarization＞68%
2.5 3He-NSF Commissioning

## Slide 31
GTAF BaF2 Array
In-situ #2
3He-NSF
6Li & 5 cm La
Neutron Detector
Un-polarized Beam
2.5 3He-NSF Commissioning
2. 3He-NSF for NOPTREX @ Back-n (2nd)
3He Polarization＞68%, Statablity~80 hours.

## Slide 32
In-situ NSF
Guide Field
Flipper
Supper Mirror
Neutron Detector
NUCL. SCI. TECH., 2022, 33:145.
Flipping Ratio vs λ
Transmission vs λ
3. 3He-NSF for Superconducting Neutron Spin Flipper
2.5 3He-NSF Commissioning

## Slide 33
4. 3He-NSF based Time-of-Flight Neutron Imaging（CSNS）
Contents | Parameters
Type | Time-of-Flight
Neutron Band | 1 – 5 Å
Flipper | In：Superconducting Spin Flipper Out：3He-NSF
Polarizing | In：Supper Mirror Out：3He-NSF
Sample | Electrified Solenoid
Parameters
Chin. Phys. Lett., 2022, 39 (6): 062901.
NUCL. SCI. TECH., 2022, 33, 145.
In-situ 3He NSF
Camera
Guide Field
Flipper
Supper Mirror
Sample
Energy Resolved
Real-space resolved
Magnetic Field
Reconstruction
P
2.5 3He-NSF Commissioning

## Slide 34
Guide field
Guide field
In-situ 3He
Camera
Neutron
3He-NSF in Front of Camera
Uniform Polarization within FOV (65%, 2.7 )
Resolution:
400 μm
FOV: 5 cm×5 cm
Deliver In-situ#2 3He-NSF from CSNS (Dongguan) to CARR (Beijing) - 2023.09
2.5 3He-NSF Commissioning
5. 3He-NSF Based Polarized Neutron Imaging at CARR (Beijing)

## Slide 35
6. 3He-NSF for Very Small Angle Scattering Spectrometer
Spin Flipper
Sample
In-situ
3He-NSF
Results of C22H43AgO2 Powder
Results for BiFeO3 SDW Spin Density Wave (SDW) Hexagonal Structure
2.5 3He-NSF Commissioning
NUCL. SCI. TECH., 2025, Accept.

## Slide 36
Outline
I. Polarized Neutrons and Hyperpolarized ³He
II. Development and Applications of ³He-NSF
III. Conclusion and Future Plans

## Slide 37
Conclusion & Plans
Fabrication of SEOP Polarized 3He-NSF (≥70%), including: Cells, Filling Station, Pumping Laser, In-situ/off-situ.
Initiated the MEOP R&D.
Conclusion
Plans
R&D 3He Cells for wide-angle neutron scattering experiments.
Scaling up MEOP system as an alternative option for 3He polarizing.
Optimizing parameters of the polarized 3He devices to quickly respond the experimental requirements from the spectrometers at CSNS.

## Slide 38: Acknowledgements
Members of Polarized Neutron Team
Group Leader: X. Tong (Tony)
Staff Members: J. Zhang, T. Wang, Z. Qin, M. Musgrave, F, Ye, L. Tian et al.
Engineer: J. Xie, Y. Li, Y, Zheng
Ph.D. Students: J. Tang, B. Wang, M. Zhang, X. Qin, Q. Zheng, T. Cui, Y. Yao, S. Wang, Q, Jiang
Master Students: G. Qiu, J. Li
Graduated Student: C. Y. Huang (Fellowship at FRM-II), Y. Dong, S. Ahmed, Y. Xiong
Thank you for your attention.
Development of polarized 3He-NSF at CSNS is a Team Effort
tongxin@ihep.ac.cn
wangtianhao@ihep.ac.cn
zhangjunpei@ihep.ac.cn
Open positions for Post Doc and Staff