## Slide 1: A (Pixel) TPC as a central tracker in a CEPC experiment
Yevgen Bilevych, Klaus Desch, Sander van Doesburg, Harry van der Graaf, Fred Hartjes, Jochen Kaminski, Peter Kluit,
Naomi van der Kolk,
Cornelis Ligtenberg,
Gerhard Raven, and
Jan Timmermans

## Slide 2: Pixel TPC
(TimePix1) TPX3 chip Quad Module TPC plane
(2007-14) 2017 2018 2019
(Octopuce)

## Slide 3
TPC requirements
The R&D on a TPC for a e.g. Linear Collider is done in the LCTPC collaboration
Requirements for a TPC from the ILC TDR
ILD like detector concept with a TPC central tracker can be used for CEPC
Challenging tracking precision
  driven by (Z) Higgs running
For Z running the requirement on d(1/pt) can be loosened by e.g. a factor 10 (LEP like)

## Slide 4: TPC layout
Material budget is
  0.01 X0 TPC gas
  0.01 X0 inner cylinder
  0.03 X0 outer cylinder
  < 0.25 X0 endplates (incl readout)
Note the very low budget in the barrel region. Material budget can be respected by different technologies like GEM, MicroMegas and Pixels
TPC is sliced between silicon detectors VTX, SIT and SET
pixel readout is a serious option for the TPC readout plane @ ILC/FFC-ee/CLIC/CEPC colliders

## Slide 5: Why a (pixel) TPC as a central tracker
Low material budget that allows high precision tracking
Continuous 3D tracking (x,y, t) and track following
  Excellent v0 reconstruction
  Ideal for particle flow and combined calorimetry
TPC is sliced between silicon detectors
  provides required momentum resolution and constraints
Very powerful particle identification based on dEdx
  for electrons (suppress pions)
  for kaons (H->ss) and protons
TPC excellent option for CEPC WW, ZH, tt etc running
Running at the Z is feasible but challenging (see discussion later)
A pixel TPC combines a large drift volume with a silicon read out plane
  tracking with high granularity and low systematic uncertainties
  due to “silicon” precision (1-10 mm) in the production process
  the best (see slides) PID performance dEdx and cluster counting
CEPC

## Slide 6: GridPix technology
Pixel chip with integrated Grid (Micromegas-like)
InGrid post-processed @ IZM (in 2025 @ Bonn)
Grid set at negative voltage (300 – 600 V) to provide gas amplification
Very small pixel size (55 µm)
detecting individual electrons
55 µm
50 µm
dyke
Aluminium grid (1 µm thick)
35 µm wide holes, 55 µm pitch
Supported by SU8 pillars 50 µm high
Grid surrounded by SU8 dyke (150 µm wide solid strip) for mechanical and HV stability

## Slide 7: Pixel chip: TimePix3
256 x 256 pixels
55 x 55 µm pitch
14.1 x 14.1 mm sensitive area
TDC with 640 MHz clock (1.56 ns)
Used in the data driven mode
  Each hit consists of the pixel address and time stamp of arrival time (ToA)
  Time over threshold (ToT) is added to register the signal amplitude
  compensation for time walk
  Trigger (for t0) added to the data stream as an additional time stamp
Power consumption
  ~1 A @ 2 V (2W) depending on hit rate
  good cooling is important
Sensitive area
2+3 mm
14.1 mm

## Slide 8: QUAD design and realization
Guard
TPX3
TPX3
TPX3
TPX3
COld
CArrier
Four-TimePix3 chips
All services (signal IO, LV power) are located under the detection surface
The area for connections was squeezed to the minimum
Very high precision 10 μm mounting of the chips and guard
QUAD has a sensitive area of 68.9%
DAQ by SPIDR system
39.6 x 28.38 mm
series of QUADs

## Slide 9: QUAD test beam in Bonn (October 2018)
ELSA: 2.5 GeV electrons
Tracks referenced by Mimosa telescope
QUAD sandwiched between Mimosa planes
  Largely improved track definition
  6 planes with 18.4 μm × 18.4 μm sized pixels
Gas: Ar/CF4/iC4H10 95/3/2 (T2K)
Ed = 400 V/cm, Vgrid = -330 V
Typical beam height above the chip: ~1 cm
Field cage
@Bonn
Published NIMA https://doi.org/10.1016/j.nima.2019.163331

## Slide 10: QUAD as a building block
in red guard wires
8-QUAD module (2x4 quads) with field cage

## Slide 11
DESY testbeam June 2021
Mounting the 8 quad module between the silicon planes
sliding it into the 1 T PCMAG solenoid
Towards a Pixel TPC part I: construction and test of a 32-chip GridPix detector, Nucl. Instrum. Meth. A 1075 (2025) 345

## Slide 12
DESY testbeam June 2021

## Slide 13
DESY testbeam Module Analysis
z global
x global
z global
y global
y local
y local
x local
z local drift
Event display with module and telescope
TPX3 track 1130 hits
c2xy = 677.5/1128
c2z = 775.9/1069
Asymmetric tail outlier removal applied 1071 hits in z kept.
TPX3 track hits
Telescope track hits (off track green)

## Slide 14
DESY testbeam Module Analysis
B=1 T p=5 and 6 GeV
T2K*
B=1 T
Fitted resolution
120 mm/
251 mm/
Ed=280 V/cm
s2 xy0= s2pixel + s2xy tele s2pixel= 552/12 mm2
sxy tele= 42 mm
Magboltz gives for DT =121 mm/
T2K* = T2K gas with O2 and H2O
ToT > 50 ms

## Slide 15
* We did not include the 4 corner chips and (11), 14, 8, 13 and 19.
These are affected by the field cage and the short in chip 11.
DESY testbeam Module Analysis
B=1T p=5 GeV
Distribution of mean residuals in the module plane
xy
z
Method row
Method column
B=1 T data set
method | rms (stat) xy | bins xy | rms (stat) z | bins z
row | 13 (2) mm | 896 | 19 (5) mm | 896
column | 11 (2) mm | 880 | 20 (5) mm | 880

## Slide 16
Tracking resolution and precision
The results of the 8 Quad Module in the DESY test beam in June 2021 have been presented
One chip (nr 11) out of 32 was disconnected due to a short*
In run 6916 e.g. 964 tracks were selected with 1009 hits on track
The tracking precision: position 9 (xy) 13 mm (z) in angle 0.19 (dx/dy) 0.25 (dzdy) mrad for a module or tracklength is 157.96 mm
The diffusion coefficients at B=0 T = 287 mm/ = 273 mm/
The diffusion coefficients at B=1 T is = 120 mm/ = 251 mm/
  In agreement with Magboltz = 121 mm/
*the chip was successfully repaired in 2023 Bonn

## Slide 17
Tracking resolution and precision
Results for the module showed that:
  the HV of the guard wires was well tuned
  B=0 T rms residuals in the module plane xy 13 mm and z 15 mm
  The results are compatible with (very) high stats quad measurement
  B= 1 T rms residuals in the plane xy 13 mm and z 20 mm;
High tracking precision is demonstrated with small systematics
  deformations in xy stay – bending plane - below 13 mm
Published in Nucl. Instrum. Meth. A 1075 (2025) 345
Towards a Pixel TPC part I: construction and test of a 32-chip GridPix detector

## Slide 18: Simulation of ILD TPC with pixel readout
To study the performance of a large pixelized TPC, the pixel readout was implemented in the full ILD DD4HEP (Geant4) simulation
Changed the existing TPC pad readout to a pixel readout
Adapted Kalman filter track reconstruction to pixels
50 GeV muon track with
pixel readout
pads
pixels
details: PhD thesis Kees Ligtenberg 2022

## Slide 19: Performance of a GridPix TPC at ILC
From full simulation the momentum resolution can be determined
Momentum resolution is about 15% better for the pixels with realistic coverage (with the quads arranged in modules coverage 59%) and deltas.

## Slide 20
Pixel TPC tracking studies
ILD tracking Performance for a Pixel TPC based on test beam
Running at B=3.5 T improves the resolution
Single electron resolution 6 mm track(“pad”) resolution
10 cm track resolution
Each 10 cm we have a point with a resolution of < 18 (31) mm on the track
Comparable to performance of a silicon detector (but TPC gas material).
T2K gas

## Slide 21
DESY testbeam Module Analysis
Particle identification (PID) performance
It is possible to study in test beam data the dE/dx or dN/dx of electrons
The Pixel TPC has measurements with 55 mm pixel size
It detects single electrons with an efficiency > 85%
This allows to measure the number of clusters (hits) as a function of the distance along the track dN/dx (dE/dx) with high granularity
The advantage of hit counting in a Pixel TPC is – due to the digital read out - that one is not including the fluctuations from the multiplication process in the charge measurement
Using e.g. a pad readout, the charge is used as a measure of dEdx
  This readout has a worse granularity and includes avalanche fluctuations
  One has to go to very small pads (‘pixels’) to measure the clusters

## Slide 22
DESY testbeam Module Analysis
Testbeam PID performance
B=0 T has a large Landau tail
B=1 T smaller Landau tail and a more gaussian distribution
An electron crossing 8 chips in the module has about 1000 hits
B=0 T
B=1 T
NIM A 1075 (2025) 345

## Slide 23
DESY testbeam Module Analysis
Analysis of PID performance
Combine chips to form a 1 m long track with 60 % coverage for electrons
Method 1 ”dEdx truncation”: reject large clusters and then run dEdx @ 90% using slices of 20 pixels along track (xy) (gives nr of selected hits). A large cluster has more than 6 hits in 5 consecutive pixels.
Method 2 “Template fit”: fit the slope of the Nscaled minimum distance (d) in xy distribution with an exponential function (Nscale(d)=defines the inverse weights):
N(d)scaled = Nscale(d) Nobserved(d)
N(d)scaled is then fitted for each track with N0 exp(-slope d)
Calculate the PID observable for electrons and MIP (==70% of hits)
  method 1 = nr of selected hits, method 2 = fitted slope
  Resolution is s = s(PID)/PID (for s we use the rms)
Published as “Towards a Pixel TPC part II: particle identification with a 32-chip GridPix detector”, Nucl. Instrum. Meth. A 1081 (2026) 170849

## Slide 24
DESY testbeam Module Analysis
PID performance method “dEdx truncation”
Electron resolution
3.6% for
1 m track and 60% coverage
Linearity MIP-e = 1.03
z drift=5-15 mm (flat)
MIP distribution is obtained by dropping 30% of the hits
T2K* gas

## Slide 25
DESY testbeam Module Analysis
B=1 T data
Calculate minimum distance between the hits.
The slope of the distribution is related to the number of primary clusters /cm
The diffused peak at d<10 (black) comes from clusters with more than 1 hit. Weights are applied at d < 10 pixels to the follow the exponential cluster distance distribution (blue). In red the exponential function.
Per 1 m track the slope is fitted to the full distance distance distribution using a ML fit.
PID performance method “Template fit”
Nucl. Instrum. Meth. A 1081 (2026) 170849

## Slide 26
DESY testbeam Module Analysis
PID performance method “Template fit”
Electron resolution
2.9% for a
1 m track and 60% coverage
Linearity MIP-e = 1.07
Ideally this is 1. A number larger than 1 means that the resolution is +7% larger

## Slide 27
DESY testbeam Module Analysis
Method | B=0 Resolution (%) | B= 1 T Resolution (%)
(1) dEdx truncation | 6.0 | 3.6
(2) Template fit | 5.4 | 2.9
The resolution for B=0 is worse than of the B=1 T data because of the larger fluctuations, that were already observed at the chip level (see slide 22).
The PID resolution for electrons from data by combining tracks to form a 1 m long track with realistic coverage ~60% coverage.
Summary of PID performance

## Slide 28
DESY testbeam Module Analysis
dEdx Performance extrapolated to the ILD detector*
Test beam B = 1 T
p=5,6 GeV/c
electron resolution 2.9(3.6)% for method 2 (1)
1 m track 60% and coverage
ILD detector
rInner = 329 rOuter = 1770 mm
halflength (z) = 2350 mm
electron resolution = 2.5(3.0)%
at q=p/2 for method 2 (1)
Assume Pixel TPC performance at B = 1 T at p = 5,6 GeV/c
*Extrapolations to CEPC were presented at the 2024 CEPC workshop

## Slide 29
DESY testbeam Module Analysis
ILD dEdx performance for T2K gas
Ullrich Einhaus performed dEdx studies in ILD and extracted the ILC soft parametrisations for energy loss based on G4 and full simulation of the ILD TPC with T2K gas
Link to the software. Samples were generated in 2020 with ILC soft v02-02 and v02-02-01

## Slide 30
DESY testbeam Module Analysis
Pixel TPC PID performance
ILD Performance with specified detector dimensions for particles at cos q = 0
Pixel TPC resolution from electron p = 5 (6) GeV test beam (for B = 1 T) of 2.5% and 3% (--- = method 1) at cos q = 0
Separation electron pion defined as:
|<Eloss e> - <Eloss p>| / s p
Separation pion kaon as:
|<Eloss p > - <Eloss K>| / s p
Nucl. Instrum. Meth. A 1081 (2026) 170849

## Slide 31
DESY testbeam Module Analysis
Pixel TPC PID performance
The expected pion-kaon separation for momenta in the range of 2.5-45 GeV/c at cos θ = 0 is more than 5.5(4.5)σ for the two resolution scenarios.
At a momentum of 100 GeV/c the separation is still 3.0(2.0)σ.
Protons can be separated from pions for momenta in the range of 2.5-100 GeV/c with more than 6.0(4.8)σ.
Nucl. Instrum. Meth. A 1081 (2026) 170849

## Slide 32
Pixel TPC PID performance
The PID resolution for an electron with p=5,6 GeV/c of 1 m track length with 60% coverage is measured to be 2.9(3.6)% at B = 1 Tesla.
The extrapolated PID resolution for an ILD detector is 2.4% (3 %)
This allows for particle identification and separation of kaons from pions up to momenta of 45 GeV with more than 5.5s (4.5s) for cos q = 0. The separation increases up to cos q = 0.85 (see back up slide).
As demonstrated, the digital read out of a pixel TPC with a small pixel size allows to perform hit and cluster counting and surpasses the PID performance of other readout schemes.
Currently a coverage of ~60% is realized with the TPX3 chip. Using the next generation TPX4 chip with Through Silicon Via‘s will allow to enlarge the coverage and increase the PID and tracking performance. The R&D is part of the DRD1 program

## Slide 33: A Pixel TPC at a circular collider
At the FCCee/CEPC a pixel TPC as a central tracker – sliced between silicon detectors as in the ILD concept detector - is well suited to carry out the WW, ZH and tt physics program.
At the CEPC a pixel TPC is selected as one of the baseline CDR/TDR detectors
A more challenging situation for a TPC is running at the Z and in particular the FCCee/CEPC Tera Z program with L = 140 1034 cm-2s-1 and Z bosons produced at a rate of 40 kHz and huge beam backgrounds*.
A 10 Giga-Z FCCee/CEPC physics program looks realistic with a GridPix pixel TPC and the current FCCee/CEPC MDI design.
*For detailed questions/answers see backup slides
Picture IHEP

## Slide 34: Conclusions: Pixel TPC at a circular collider
The performance of a pixel TPC based on GridPixes has been studied based on prototypes and test beam measurements. It provides
  high precision continuous tracking with a low material budget
  very powerful particle identification
At the CEPC a pixel TPC as a central tracker – sliced between silicon detectors as in the ILD concept detector - is well suited to carry out the WW, ZH and tt physics program.
A 10 Giga-Z physics program looks realistic with the proposed pixel TPC and the current MDI design.
After years of R&D, a pixel TPC has become a realistic viable option for experiments
R&D on a Pixel TPC with GridPix devices with reduced ion back flow based on the TPX3 or TPX4 ASICs are part of the DRD1 program.

## Slide 35: Backup plots

## Slide 36
Pixel TPC tracking studies
ILD tracking Performance for a Pixel TPC based on test beam
The last 10 cm track provides very high resolution ‘point’ in the endcap (cos q>0.8). This is due to the short drift distance and the high resolution pixel readout.
This high precision endcap ‘point’ can be used to calibrate out the TPC distortions presents at high luminosity Z running.

## Slide 37
DESY testbeam Module Analysis
Pixel TPC dEdx performance
Separation pion kaon
|<Eloss p > - <Eloss K>| / s p
Separation pion kaon for different cos(theta) values due to the track length dependence
For cos(theta)=0 till 0.95 the separation lies between the black and red curves. Only above 0.95-0.975 the separation drops till the blue curve.
Excellent performance over very large polar angle range
ILD Pixel TPC

## Slide 38: A Pixel TPC at a circular collider
Can a pixel TPC reconstruct the Z events?
  The TPC total drift time is about 30 μs
  This means that there is on average 1.2 event / TPC readout cycle
  YES: The excellent time resolution: time stamping of tracks < 1.2 ns allows to resolve and reconstruct the events
Can the current readout deal with the Z rate?
  Link speed of Timepix3 (in Quad): 2.6 M hits/s per 1.41 × 1.41 cm2. In the module testbeam we tested up to rates of 5.7 kHz.
  YES: This is sufficient to deal with hits from Z’s in high luminosity Z running
    Expect about 90 kHz/cm2 from hadronic Z decays@inner radius
What is the current power consumption?
  No power pulsing is possible at a circular collider (at e.g. ILC power pulsing is possible)
  Current power consumption TPX3 chip 2W/chip per 1.41 × 1.41 cm2
  So: good cooling is important but no show-stopper.
  In LCTPC, CO2 cooling has been developed and tested for the MM read-out technology
  To save power the TPX3/4 chips can be run in LowPowerMode: reduction factor 10.

## Slide 39: A Pixel TPC at a circular collider
The importance of the beam background and the optimization of the FCCee/CEPC machine detector interface MDI
  Hit rates are currently dominated by the beam background (not by Z decays)
  The beam background study presented in the October 2024 ECFA meeting by Daniel Jeans (for a 2 T B field) would correspond to a sizeable background rate of about 300 MHz/cm2 (@inner radius).
  Since October 2024, the MDI has improved, a B field of 2.5(3-3.5?) T seems possible and the backgrounds have been reduced significantly. But his needs more optimisation.
  The Pixel TPC – due to the high granularity – has an occupancy of < 1% for 300 MHz/cm2 (@inner radius). The read-out speed of the current TPX3 has to be upgraded to cope with the large background rate. R&D for a fast read-out based on TPX3/4, like the VeloPix (LHCb) is needed.

## Slide 40: A Pixel TPC at a circular collider
What about track distortions?
  Not an issue for running at cms energies of WW, ZH, tt
  There are two important sources of track distortions of the TPC drift field:
    the slowly drifting primary ions (from Zs and background)
    the ions produced in the amplification process flowing back (IBF)
  At the ILC gating is possible to reduce the IBF; at FCC-ee this is not realistic, for a Pixel TPC a double grid is the best solution (see below)
Is it possible to reduce the IBF for a pixel TPC?
  IDEA: by making chip with a double grid structure (see back up slide)
  This idea was already realized as a ‘TWINGRID’ NIMA 610 (2009) 644-648
  For GEMs for the ALICE TPC this was also the way – use several GEMs to reduce IBF
  The IBF* can be easily modelled and with a hole size of 25 μm an IBF of 3 10-4 can be achieved and the value for IBF*Gain (2000) would be 0.6.
  YES: the IBF can be reduced. It needs R&D that can be done in the detector lab in Bonn.
  * Measured IBF of single grid GrixPix is 1.3% and IBF*Gain (2000) is 26.

## Slide 41: A Pixel TPC at a circular collider
What would be the size of the TPC distortions at a TeraZ factory?
  Tera-Z studies by Daniel Jeans and Keisuke Fuji show that for FCC-ee this means: distortions from Z decays up to < O(100) mm
  The problem is that the beam backgrounds are by far the dominant source for the ions in the TPC volume. A reduction of that background is important.
  The beam background study presented in the October 2024 ECFA meeting by Daniel Jeans (for a 2 T B field) gave estimates for the old MDI but it is likely that the numbers can go down substantially. The MDI for ILC (ILD_I5_v02_2T) yields a background that is a factor 40 lower than the FCCee.
  There is also an idea to operate the TPC with a Neon based gas that could bring the backgrounds down by an order of magnitude. This needs study.
  It can be argued that in an ILD like detector the distortions can be mapped out using the silicon VTX-SIT/SET detectors. See e.g. the backup slide on fitting out TPC distortions.

## Slide 42: A Pixel TPC at a circular collider
A Tera Z CEPC physics program* could be carried out, but
  It needs a lot of power to process the large amounts of background hits in the tracking detectors
  It needs more study to reduce the beam background by an improved MDI
  Fitting techniques should be studied to correct for the TPC distortions
  The current TPX3 read out should be optimized to cope with the high readout rate and minimise the power consumption
* For the CEPC the Tera Z option is considered to be less important. The physics case seems IMO rather weak to go beyond 10 Giga Zs.

## Slide 43: Fitting out TPC distortions
It is possible to map out distortions using e.g. muons from Z decays
  E.g. by fitting the 3D spatial distribution as a function of time as was done by ALEPH and more recently by ALICE. Using this distribution the hits positions are corrected and the TPC track refitted.
However, with silicon trackers around the TPC, more elaborate methods can be used. One can use the track predictions based of the silicon trackers SIT and SET to correct on a track-by-track level the TPC track.
  One can use as a constraint that the extrapolated positions and angles agree with the measured in the SIT and SET.
  Practically, one can e.g. correct the TPC track parameters
The ultimate way is a fitting technique similar to ATLAS. In the ATLAS track fit the common systematics is fitted out for sets of Muon hits. For ILD @ FCC the fit would fit free parameters in the distortion model, while using as a constraint the SIT and SET position and direction measurements.
  The simplest case is a model where the strength (amplitude) and radial dependence would be scaled and a model is used for the 3D extrapolations.

## Slide 44: Reducing Ion back flow in a Pixel TPC
The Ion back flow can be reduced by adding a second grid to the device.
It is important that the holes of the grids are aligned. The Ion back flow is a function of the geometry and electric fields. Detailed simulations – validated by data - have been presented in LCTPC WP #326.
With a hole size of 25 μm an IBF of 3 10-4 can be achieved and the value for IBF*Gain (2000) would be 0.6.
Ion backflow | Hole 30 μm | Hole 25 μm | Hole 20 μm
Top grid | 2.2% | 1.2% | 0.7%
GridPix | 5.5% | 2.8% | 1.7%
Total | 12 10-4 | 3 10-4 | 1 10-4
transparancy | 100% | 99.4% | 91.7%