## Slide 1
Xiyang Wang Shiming Zou
2024.3.20
SiPM PreAMP for ASHIPH（updata）

## Slide 2
Low Temperature System
semiconductor chilling plate
PCB
MCX
SiPM
AMP
Temp Sensor
Cold Water
Flow
16CH ADC
Temperature monitoring
Temp Sensor
Room Temperature
8CH Digitalizer
DT5725
MON CH1
MON CH0
USB
USB

## Slide 3
Low Temperature System
Start cooling
Temperature begins to stabilize：- 3 ℃

## Slide 4
Dark counts test (For 5SiPM Array )
LOW
Temperature
~-3℃
Vbias = 31V
Vbias = 32V
Vbias = 33V
Room
Temperature
~25℃
Vbias = 31V
Vbias = 32V
Vbias = 33V

## Slide 5
Dark counts test (For 5SiPM Array)
Peak 1
Peak 2
Peak 3
Least Squares
SPA(Single photon amplitude) = b
SNR(Signal to Noise Ratio) = b/σ1
SPA VS Bias
SNR VS Bias

## Slide 6
Performance of different PCB

## Slide 7
Performance of different PCB
Multiple peaks due to inconsistent SiPM performance

## Slide 8
Performance of different PCB
Board NO. | 1 |  | 2 |  | 3 |  | 4 |  | 5 | 
Temperature | 25℃ | -3℃ | 25℃ | -3℃ | 25℃ | -3℃ | 25℃ | -3℃ | 25℃ | -3℃
SPA（ADC） | 214.9 | 294.8 | 240.4 | 317.8 | 222.9 | 304.4 | 311.3 | 384.3 | 246.7 | 321.9
SNR | 7.4 | 10.5 | 7.1 | 9.4 | / | / | 9.4 | 12.6 | 7.7 | 10.5
Dark Current(uA) | 2.52 | 0.73 | 1.47 | 0.52 | 1.47 | 0.52 | 2.06 | 0.71 | 9.55 | 8.55

## Slide 9
Dark counts test (For 15SiPM Array )
LOW Temperature
~-3℃
Room Temperature
~25℃
SPA(Single photon amplitude) = 178.1
SNR(Signal to Noise Ratio) = SPA/σ1=5.2
SPA(Single photon amplitude) = 226.5
SNR(Signal to Noise Ratio) = 7.1

## Slide 10
Back Up

## Slide 11
Back UP data

## Slide 12
Back UP data

## Slide 13
Back UP data