## Slide 1
More than 10 Years Struck mTCA Modules for LLRFcontrol and Beam Instrumentation
muTCA/ATCA product page (struck.de)
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH
RONG LIU

## Slide 2: Outline
Digitizers
Mid Speed High Resolution (125/250 MSPS 16-bit) Data Streaming
LLRF BPM
High Speed High Resolution (2.5 GSPS 14-bit) FCT
Picoammeter Spin off (1MSPS 20-bit)
Digital I/O SIS8864 I/O
SIS8800 Multi Purpose Scaler
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 3: SIS8300-KU 10 channel 125 MSPS 16-bit digitizer (250 MSPS 14-bit stuffing option)
History
Virtex 5 based
2010 SIS8300 initial development, first MTCA.4 board worldwide
2011 SIS8300 V2 > 370 units in field (in part 8 channel 250 MSPS 14-bit BPM)
Virtex 6 based
2014 SIS8300-L > 150 units in field
2015 SIS8300-L2x > 930 units in field (L2S, L2D,…)
SIS8300
SIS8300-L
SIS8300-L2
SIS8300-KU
SIS8325
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 4: SIS8300-KU Properties
10 Channels 125 MS/s 16-bit or 250 MS/s 14-bit ADC
10 MS/s to 125 MS/s Per Channel Sampling Speed
AC or DC Input Stage
Internal, Front Panel, RTM and Backplane Clock Sources
Two 16-bit 250 MS/s DACs for Fast Feedback Implementation
High Precision Clock Distribution Circuitry
Programmable Delay of Dual Channel Digitizer Groups
Multi Gigabit Link Port Implementation to Backplane
Twin SFP+ Card Cage for High Speed System Interconnects
White Rabbit Clock Option for SFP+ Ports
Two RJ45 Connectors (One Clock + 3 Data or 4 Data In/Out)
XCKU040-1FFVA1156C Kintex Ultrascale FPGA
2 GByte DDR4 Memory (flexible partitioning scheme)
4 lane PCI Express Gen3 Connectivity
Dual boot
MMC1.0 under DESY license LV91
In Field Firmware Upgrade Support
Vivado Project for Custom Firmware Development
Zone 3 class A1.0, A1.0C or A1.1CO Compatible
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 5: SIS8300-KU White Rabbit (WR) Ready
PLL Clock Synthesizer
VCO, VCXO
DACs for VCO Control
Additional Power
CERN Kintex Ultrascale White Rabbit PTP Core First DESY Techlab Tests
What is WR?
Sub ns synchronisation Synchronous Ethernet IEEE 1588 PTP
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 6: FPGA Resources/Performance
Used FPGA | XC6VLX130T-2FFG1156C (SIS8300-L2) | XCKU040-1FFVA1156C (SIS8300-KU)
System Logic Cells | 128 K | 500 K
CLB LUTS/Flip-Flops | 80/160 K | 242/484 K
Block RAM/FIFO w/ECC (36Kb each) | 264 | 600
Block RAM/FIFO (18Kb each) | 528 | 1200
Total Block RAM (Mb) | 9.5 | 21.1
DSP Slices | 480 | 1920
GTH speed | 6.6 Gb/s | 16.3 Gb/s
PCIe | Gen2 | Gen3
Memory Controller | DDR3 | DDR4
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 7: SIS8300-KU AXI Based VIVADO Project
Cooperation Lund University/ESS/Struck
To Meet ESS/ERIC Firmware Specs
ESS Version → COSYLAB EPICS
Struck Version available to all users
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 8
System Integration Examples
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 9
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 10: SIS8300-x ROOT GUI SIS8300-KU 8AC2DC with SIS8900 RTM
RMS for open AC/DC channel
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 11: SIS8300-KU Wendelstein X7 Application 40 Channel Streaming System
empty | empty | empty | SIS8900 | SIS8900 | SIS8900 | SIS8900 | empty | empty
AC600D | MCH | empty | SIS8300-KU | SIS8300-KU | SIS8300-KU | SIS8300-KU | empty | empty
Rear
Card Cage
Front
Card Cage
Native R5 MTCA.4 Crate
MCH
NAT PHYS80
8 Lane optical PCIe Uplink
16 Lane future option
8 Lane PCIe
Uplink
DAQ PC/Workstation
Uplink
8 Lane optical PCIe Uplink
16 Lane future option
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 12: SIS8300-KU Application Example 40 Channel Streaming System (Front)
8 Lane PCIe Uplink
NAT PHYS80
4 x SIS8300-KU
PCIe Uplink Led
External Sample Clock
Input Voltage: 0.3V to 5V
Input Frequency: 1 MHz to 50 MHz
External Trigger Input
RJ45 LVDS Input 1
PCIe Link Led
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 13: SIS8300-KU System Example 40 Channel Streaming System (Rear)
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 14: Fast Langmuir Probes on Wendelstein 7-X
Wendelstein 7-X is a nuclear fusion experiment in Greifswald, Germany at the Max-Planck-Institute for Plasma Physics
Langmuir probes measure plasma density and temperature
Fast probes can resolve turbulence
40 ADC channels with 16b at 10 MS/s
Synchronised with
~50 other systems
In operation for
~500 experiments of up to 100s
Courtesy Lukas Rudischauser, IPP
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 15: SIS8300-x/DWC8VM1 as single cavity LLRF solution covering 350 MHz to 6 GHz
Current combination: SIS8300-KU Kintex Ultrascale Digitizer and DWC8VM1
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 16: DWC8VM1 8 Channel Downconverter One Channel Vectormodulator*
*under license from DESY
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 17: DWC8VM1 Properties
MTCA.4 RTM Implementation
8 Channels Downconverter
350 MHz - 6 GHz (3 different types)
8 Channel FBM Multi Coax. Connector (CH1 to CH8)
2 Auxilliary Channels
One Channel Vector Modulator
VM Output 50 MHz to 6 GHz
SMA Vector Modulator Output
Various Intermediate Frequencies
Switchable Front End Attenuators
LO Clock From Front Panel or RF Backplane
LO Power Level Monitor
Interlock Scheme
I2C Support
Zone 3 Class A1.1 compatible
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 18: DS8VM1 8 Channel Direct Sampling One channel vectormodulator*
*under license from DESY
SIS8300-x/DS8VM1
as single cavity LLRF solution covering 5 MHz to 700 MHz Undersampling in some cases
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 19: DS8VM1 GSI 108 MHz Input Filter and Input Response
Jens Zappai et. al. LLRF2017
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 20: SIS8300-x BPM Application
PAL FEL
144 Stripline BPMs SIS8300 V2
in 8 channel 250 MSPS configuration SLAC RTM
56 Cavity BPMs SIS8300-L2/SLAC RTM DESY FLASH
SIS8300-L2/DESY Interleaving RTM ESS
SIS8300-KU/DWC10LF
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 21
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 22: SIS8325 250 MSPS 16-bit Digitizer (Helmholtz Validation fund project)
MTCA.4
4 lane PCI Express Connectivity
10 Channels 250 MS/s 16-bit ADC
10 MS/s to 250 MS/s Per Channel Sampling Speed
AC and DC Input Stage
Internal, Front Panel, RTM and Backplane Clock Sources
Two 16-bit DACs for Fast Feedback Implementation
High Precision Clock Distribution Circuitry
Programmable Delay of Dual Channel Digitizer Groups
Gigabit Link Port Implemenation to Backplane
Twin SFP Card Cage for High Speed System Interconnects
Virtex 6 FPGA
One GSample Memory (flexible partitioning scheme)
In Field Firmware Upgrade Support
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 23: Application FAIR FCT (Fast Current Transformer)
Test System with SIS3350 500 MSPS 12- bit VME Digitizer
Development of new MTCA FMC Carrier and 2.5 GSPS 14-bit Digitizer FMC
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 24: SIS8160 Dual FMC Carrier AMC
4-lane PCI Express Gen3 Connectivity
Xilinx XCKU40- or XCKU060- 1FFVA1156C Kintex Ultrascale FPGA
Dual Boot
Front Panel MMCX Clock Input
Front Panel MMCX Digital In-/Output (HW Configuration)
Point to Point Links
4 MLVDS µTCA Ports (AMC Ports 17-
20) → 8 MLVDS lines
2 HPC FMC Sites
Variable FMC VADJ (1,0V - 1,8V)
Low Jitter Clock Generation and Management
2 x 2 GByte DDR4 Memory with two Memory Controllers
White Rabbit Option (over FMC 2)
Stand Alone Operation Option
MMC1.0 under DESY LV91
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 25: SIS8160 Block Diagram
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 26: SFMC01 2 Channel 2.5 GSPS 14-Bit
Single width, 10mm stacking height, air cooled commercial grade HPC FMC Module
Analog front-end factory configurable for DC- or AC (Balun) input coupling
Up to 5 GHz AC Analog BW
Up to 1 GHz DC Analog BW
Dual channel 14-Bit, 2.5 GSPS with JESD204B Interface
50 Ohm Input Termination
4 Front panel SMA Inputs for Analog A/B, Clock and Trigger
Ultra low phase noise 100 MHz on board clock source
High performance jitter attenuating frequency generator for JESD204B
one green Front panel user LED
System management EEPROM and Temperature Sensor with Thermal Watchdog
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 27: SFMC01 2 Channel 2.5 GSPS 14-Bit
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 28: SIS8160/SFMC01 RROT GUI DC Configuration 1 GHz BW
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 29: SFMC01 AC 5 GHz BW Open Input
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 30: SFMC01 Gain Flatness request: +/- 0,5 dB 0-100 MHz
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH
Notes: Possible next FMC JESD Developments
4 Channel GSPS 16-bit Digitizer
8 Channel 500 MSPS 14-bit Digitizer
Project Driven …
ADC/DAC Combination …

## Slide 31
FCT is used for optimization of the „ bunch merging 4 → 2“ process: cancelling bunch oscillations | Measurement 1 Measurement 2
Waterfall view
Contour view
operational at SIS18 and ESR
state-of-the-art DAQ based on µTCA 2.5 GSa/s, 14-bit ADC
rf-triggered multi-event acquisition (programmable rate divider for long acqusition times in development by BEA)
DAQ and operator GUI by BEA
SIS18 FCT
Initial Application: FAIR Fast Current Transformer (FCT) System
Front End Electronics
l
l
l
l
l
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 32: SIS8160 Application CAENELS FMC-Pico-1M4(-C3)
4 Channel Pico Ammeter
20-Bit 1 MSPS (BW 10/C3 300 KHz)
RNG0: ± 1 mA, RNG1: ± 1 µA
C3: RNG0: ± 10 mA, RNG1: ± 500 µA
BPM, SR Beamline,…
SIS8300-KU like FW framework
Vivado upgrade path to older carrier HW
Picture courtesy CAENELS
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 33: SIS8160/FMC Pico Noise Distribution (reminder 20-bit)
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 34: SIS8864 64 channel LVTTL Digital I/O AMC
Machine cycle dependent & generic ouput pattern with digital input
Alternative to OHWR FMC carrier and two I/O FMC‘s
Ready to run firmware
Application specific firmware adaptations
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 35: SIS8864 64 channel LVTTL Digital I/O AMC Properties
AMC with Double Width Mid-Size form factor
Xilinx XC7A15T-2FGG484C Artix-7 FPGA
Single lane PCI Express Gen2 Interface
1 AMC Port GbE
2 AMC Ports Point-to-Point Serial Link
4 AMC Ports MLVDS (8 MLVDS lines)
2 Front panel 32 data I/O: Mini D Ribbon (MDR) (TTL/LVTTL)
I/O direction programmable in 8-Bit Groups
1 Front panel control Input: LEMO (TTL/LVTTL)
1 Front panel control Output: LEMO (LVTTL)
In field firmware upgrade capability
Module Management Controller ATxmega128A1U, IPMB-L interface
DESY MMC1.0 (under LV 91)
In production
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 36: Typical FAIR SIS8864 Digital I/O Use Case
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 37: SIS8864 Digital I/O Firmware Detail Event Receiver/MLVDS → Pattern Memory
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 38: SIS8800 Multi Purpose Scaler
Functionality
MTCA.4 AMC
4 Lane PCI Express connectivity
XC6VLX130T-2FFG1156C Xilinx
Dual boot
Redundant PCIe implementation
2 GByte DDR3 memory
16 front inputs NIM or TTL/LEMO, TTL,ECL or LVDS/flat cable
200 MHz count rate (NIM/ECL)
4 control in-/4 control front outputs
Zone 3 Class D1.1 compatible
MMC 1.0 under DESY license LV91
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 39: SIS8890 Discriminator RTM for SIS8800 Multi Purpose Scaler
Functionality
MTCA.4 RTM
16 Discriminator Inputs, Connector Type MMCX
Leading Edge
50 Ω Input Termination
14 Bit Threshold DAC for each Channel, 0.4 mV/Step
16 Outputs to Zone 3 (to Scalers on SIS8800)
16 Front Outputs, Connector Type MMCX
Programmable Pulse Width 10 ns -250
Zone 3 Class D1.1 Compatible
MMC1.0 Compatible
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 40: SIS8980 in NIM/MMCX Configuration
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 41: MTCA.4 Relevance, current Struck user base
AU | Australian Synchrotron
BR | LNLS
CH | CERN
CN | IHEP, SINAP, USTC, CIAE,CSNS,IMP,DLS,IASF,PKU,THU,CQU,…
CZ | ELI (Inst of Physics, Praha)
DE | DESY, HZDR, PTB, MPG, KIT, HZB, GSI, DESY Zeuthen, HMI
ES | ESS Bilbao, GMV
FR | ITER, Saclay
GB | Diamond, STFC
IN | TIFR
JP | KEK, SPring-8
KR | PAL
RU | ITER, JINR (Techlab/Bevatech)
SE | ESS, Lund University
TR | TARLA
TW | NSRRC
US | SLAC, NSCL/FRIB, ANL, ORNL
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH

## Slide 42: Conclusion/QA/&thanks
MTCA(.4) standard widely adopted, Companies and DESY as incubators
Increasing number of users and mtca based designs and solutions in China now
Looking forward to more exchanges and discussions to promote the development of this technology and the community
partner
Beijing DAQ Technology Co. Ltd.
The 2025 MTCA/ATCA for Large Scientific Facility Control Workshop, CQU Chongqing, 17.09.2025
© 2025 Copyright Struck Innovative Systeme GmbH