## Slide 1
Low-Level Radio Frequency Control System
for HUST-UED
Reporter: Changda Peng
Huazhong University of Science and Technology

## Slide 2
Experimental Results
Laser-RF Synchronization
Development of LLRF System
HUST-UED RF Requirements
2
1
3
4
Summary
5
Notes: 接下来我将分为5个部分进行汇报

## Slide 3
HUST-UED RF Requirements
1
Notes: 接下来我将分为5个部分进行汇报

## Slide 4
Ultrafast Electron Diffraction Technique
Atomic motion in matter ≈ 100 fs
Temporal resolution for UED facility < 100 fs
Femtosecond laser pump
Ultrashort electron probe
Diffraction patterns
Sub-nanometer-scale
Spatial Resolution
Femtosecond
Timescale Resolution
τprobe | Temporal length of the electron probe
τpump | Temporal length of the pump laser
τvm | Velocity mismatch-induced temporal jitter
τjitter | Electron beam arrival temporal jitter
Pump-Probe Technique
Notes: Ultrafast electron diffraction (UED) employs a pump-probe strategy, where a femtosecond laser pump transiently excites the sample and an ultrashort electron beam serves as the probe to capture instantaneous diffraction patterns.
This approach simultaneously achieves femtosecond-level temporal resolution and sub-nanometer-scale spatial resolution.
The characteristic timescale of atomic motion in matter is approximately 100 femtoseconds. Therefore, to accurately observe the dynamic processes of structural changes, the temporal resolution of a UED setup must be better than 100 fs.

## Slide 5
HUST-UED Pump-probe Experiment
Amplitude and phase of the RF electromagnetic field are controlled by LLRF control system
UV laser(266 nm)
Photocathode
Electron bunch
Beam energy 3 MeV
1.4-cell RF gun
1.4-cell RF gun:
Acceleration field gradient 76.6 MV/m
Optimal injection phase 65°
Resonant frequency 2856MHz
Suppress space charge effects and compressing the arrival temporal jitter
Notes: UV laser irradiation of the photocathode
The high-gradient electric field rapidly accelerates the electron beam to relativistic speeds, suppressing the space charge effect

## Slide 6
HUST-UED RF System
Adjust the delay between signals to enable the low-level chassis to synchronously sample the RF power
Klystron
Microwave
Power Source
Waveguide
RF Power Transmission
Electron Gun
Generate
electron beam
Phase shifter compensates
for phase shifts
Directional couplers to
monitor RF power
Notes: The microwave power is generated by a klystron and transmitted to the electron gun via a waveguide..

## Slide 7
2
Development of LLRF System
Notes: 接下来我将分为5个部分进行汇报

## Slide 8
HUST-UED LLRF Control Platform
AMC and RTM enable the acquisition, processing, and feedback control of RF signals
MTCA.4 architecture
Notes: Based on requirements and simulations, we have developed an LLRF control platform.

## Slide 9
HUST-UED LLRF Control Platform
AMC and RTM are combined via the data transmission interface Zone3
The output signal frequency is 2856 MHz, consistent with the operating frequency of the RF system
Notes: Based on requirements and simulations, we have developed an LLRF control platform.

## Slide 10
HUST-UED LLRF Control Platform
Digital signal processing module :
High-speed signal acquisition, Amplitude and phase calculation, Feedback control
RF front-end module:
Down-conversion of RF signals, Output signal conditioning, Signal attenuation
Notes: Based on requirements and simulations, we have developed an LLRF control platform.

## Slide 11
Signal Processing Algorithm
Mix RF Signal with LO Signal
Adjust the vector angle based on the received Φ
Down-conversion
IQ Modulation
Decompose the signal into I and Q components
PI
Vector Rotation
Vector Modulator
Conversion from I/Q signals to a sinusoidal signal
Generate correction value
CORDIC
Clockwise rotation of the IQ vector
Notes: Based on requirements and simulations, we have developed an LLRF control platform.

## Slide 12
Two-Stage PI Closed-Loop Control
"Fast Loop" running on FPGA
Direct I/Q vector control for real-time RF field stabilization
"Slow Loop" running on CPU
Addressing RF Cavity Detuning
Notes: The amplitude and phase stability of the VM signal generated by the LLRF system are relatively poor in open-loop control mode, as it is susceptible to environmental factors (e.g., temperature, humidity) and internal equipment noise.
We introduced drift and jitter in the simulation to emulate real-world conditions.
The resonant frequency drift is converted into a phase drift slope, which is computed by the CPU to adjust the water-cooling temperature for fine-tuning the cavity dimensions, thereby achieving resonant frequency tunning.

## Slide 13
Two-Stage PI Closed-Loop Control
"Fast Loop" running on FPGA
Direct I/Q vector control for real-time RF field stabilization
"Slow Loop" running on CPU
Address RF Cavity Detuning
The thread sleeps for 2 seconds after completing a full measurement and control cycle
Notes: The amplitude and phase stability of the VM signal generated by the LLRF system are relatively poor in open-loop control mode, as it is susceptible to environmental factors (e.g., temperature, humidity) and internal equipment noise.
We introduced drift and jitter in the simulation to emulate real-world conditions.
The resonant frequency drift is converted into a phase drift slope, which is computed by the CPU to adjust the water-cooling temperature for fine-tuning the cavity dimensions, thereby achieving resonant frequency tunning.

## Slide 14
Signal Attenuation Calibration and Matching
Digital attenuation limits all channels to within input range
Real-time attenuation compensation with embedded algorithms
Logarithmic detector
Provides accurate mathematical basis for real-time power feedback and closed-loop control
Digital attenuation
Attenuation compensation
Digital attenuation
Notes: The three-stage attenuation link, composed of a directional coupler, RF Cable, and attenuator, acquires RF signals.

## Slide 15
3
Laser-RF Synchronization
Notes: 接下来我将分为5个部分进行汇报

## Slide 16
HUST-UED Laser-RF Synchronization
Laser Phase
RF Phase
DG645
1/36 Frequency
Divider
Signal Source
Low-level chassis
Solid-State
Amplifier
RF Front-End
Clock & local oscillator
Trigger Transition
LLRF System
Laser System
RF System
Laser-RF System Signal Connection
HUST-UED LLRF Rack
Notes: The laser pulse must be synchronized to irradiate the photocathode precisely when the RF field in the accelerating cavity reaches the optimal injection phase. This enables the electron beam to achieve higher acceleration energy while simultaneously compressing its time-of-flight (TOF) to nearly the minimum value.

## Slide 17
HUST-UED Laser-RF Synchronization
Laser-RF System Signal Connection
Laser System Synchronization
Laser Internal Oscillator Signal
1/36 Frequency Division Signal
Use photodiode to detect the frequency and phase difference
Notes: A 2856 MHz signal from the master oscillator's power divider output is divided by 36 to 79.3 MHz, which is then used for phase-locking with the Vitara seed source inside the laser system. This signal achieves precise synchronization with the Ti:sapphire oscillator's optical pulses through a three-stage closed-loop system.

## Slide 18
HUST-UED Laser-RF Synchronization
Laser-RF System Signal Connection
LLRF System Synchronization
Trigger Pulse
SSA Output
Pulse width strictly follows that of the trigger signal
Notes: A 2856 MHz signal from the master oscillator's power divider output is divided by 36 to 79.3 MHz, which is then used for phase-locking with the Vitara seed source inside the laser system. This signal achieves precise synchronization with the Ti:sapphire oscillator's optical pulses through a three-stage closed-loop system.

## Slide 19
HUST-UED Laser-RF Synchronization
Laser-RF System Signal Connection
Laser-RF Synchronization
RF System Synchronization
Laser
RF gun
Modulator
Klystron
Circulator
RF gun input
Meet the waveform requirements for the input RF power
Notes: The pulse width of the output trigger signal from the DG645 delay generator is adjusted.
The output signal waveform precisely follows variations in the trigger pulse width while maintaining high temporal synchronization.

## Slide 20
4
Experimental Results
Notes: 接下来我将分为5个部分进行汇报

## Slide 21
Low-Level Chassis Test Results
Input Test
Output Test
Input Signal
LO Signal
REF Signal
CLK Signal
Input signals of the low-level chassis meet both the frequency requirements and the power constraints
Output signal frequency of the low-level chassis is 2856 MHz, maintaining consistency with the REF signal.
RMS Amplitude Stability
0.0513 %
RMS Phase Stability
0.0375°

## Slide 22
HUST-UED Experimental Test Results
Amplitude and Phase Waveform Detection
Test Signal RMS Stability
Test signal | RMS Amplitude Stability | RMS Phase Stability
VM | 0.0513 % | 0.0375°
Klystron | 0.4866 % | 0.2981°
Open-Loop Input | 0.5007% | 0.2806°
Close-Loop Input | 0.0895 % | 0.1349°
Closed-loop control significantly improves RF power stability, thus stabilizing the accelerating field
Open-Loop
Close-Loop
Notes: The fed power is largely absorbed by the electron gun and is subsequently released after the feeding process is completed.

## Slide 23
HUST-UED Experimental Test Results
Amplitude and Phase Waveform Detection
Test Signal RMS Stability
As the input power increases, the RMS signal stability improves
Amp Stability
Phase Stability
Notes: Root Mean Square

## Slide 24
HUST-UED Experimental Test Results
Amplitude and Phase Waveform Detection
Dark Current Observation
The electron beam focused by the solenoid coil can be observed on the fluorescent screen
Main Current Observation

## Slide 25
5
Summary
Notes: 接下来我将分为5个部分进行汇报

## Slide 26
Summary
High-precision and high-stability LLRF control system developed for HUST-UED
LLRF system employs an MTCA architecture
Integrate RF front-end and digital processing modules
Enable RF signal acquisition, processing, and closed-loop feedback control
Output signal amplitude and phase exhibit high stability
Meet electron beam commissioning requirements！

## Slide 27
THANKS FOR YOUR LISTENING