## Slide 1: Investigation of Diagonal Cut-Plane BPM Performance in the CSNS RCS
M. A. Rehman, Xu Zhihong, Qiu Ruiyang, Renjun Yang, Yang Rui
2025-05-06

## Slide 2: Content
Reveal the Edge Effect in RCS BPM
  The beam pipe and BPM cross sections/shapes are different
  The electromagnetic field of the beam is distorted in the vicinity of the edges of the BPM
  Analytical results, first order [1]
  J-PARC MR also reported a similar issue [2, 3]
  We use CST simulation to reveal the edge effect
Calibration of the BPM with wire mapping
  Alignment of BPM with wire
  Mover precision
  Mapping precision (17 times repeated)
  Mapping with different frequencies
[1] J. H. Cupérus, NIM145 (1977) 233-243
[2] T. Toyama, IPAC10 MOPE012
[3] ] T. Toyama, OHO School 2009 (Japanese)

## Slide 3: Edge Effect in BPM
The beam pipe and BPM cross sections/shapes are different
The electromagnetic field of the beam is distorted in the vicinity of the edges of the BPM
  Design a BPM in such a way as to prevent the edge effect
    Guard rings in an upstream and downstream direction to prevent the edge effect
    The longitudinal space is limited!
  Wire Mapping
    The wire is a constant-potential surface
    For the beam, the potential along its length depends on the charge of the beam and the cross-section of the beam pipe
  Mapping with Beam [only one such setup at ANL ZGS (J. R. Simanton, 1969)]!
    Expensive and difficult to realize the installation
  Analytical formulae
    Less accurate, valid only for simple geometries
  FEM simulations CST
    Wake field solver line charge only few weeks!
    PIC solver beam transverse cross-section few months
[1] J. H. Cupérus, NIM145 (1977) 233-243
[2] T. Toyama, IPAC10 MOPE012

## Slide 4: CST Simulations
Only BPM
  Bunch charge 1 µC, Bunch Length 3000 mm/1σ (10 ns/1σ), β = 1
  ¼ Region, 0 to x/y:60 mm, with 10 mm step
  Bunch charge 1 µC, Bunch Length 7000 mm/1σ (24 ns/1σ ), β = 1
    ¼ Region, 0 to x/y:60 mm, with 10 mm step
BPM with Upstream Bending Magnet Pipe
  Bunch charge 1 µC, Bunch Length 3000 mm, β = 1
    ¼ Region, 0 to x:60 mm, y: 40 mm, with 10 mm step
Realistic: BPM with Upstream Bellows and Bending Magnet Pipe
  Bunch charge 1 µC, Bunch Length 7000 mm, β = 1
    ¼ Region, 0 to x:60 mm, y: 40 mm, with 10 mm step
Wire Calibration
  Mapping Data 1 MHz, Step size 1 mm, 17× Repeat
  Mapping Data (0.5, 1, 1.5, 2.0, 2.5, 3, 4, 5, 6, 7, 8, 9, 10) MHz, Step size 1 mm, Repeat 1, W/ Temp. Monitoring

## Slide 5
C11 C12 C13 C14
C21 C22 C23 C24
C31 C32 C33 C34
C41 C42 C43 C44
(
(
=
296.62 pF -8.56 pF -3.07 pF -1.40 pF
-8.56 pF 300.1 pF -7.50 pF -3.07 pF
-3.07 pF -7.50 pF 300.44 pF -8.54 pF
-1.40 pF -3.07 pF -8.54 pF 296.60 pF
220 mm
226 mm
CST Simulations: Capacitance and S-Parameters
Electrode Isolation
RCS BPM Model
Capacitance Matrix

## Slide 6: CST Simulated Models

## Slide 7
220 mm
226 mm
83.96 mm
24 mm
77 mm
135 mm
218 mm
3 mm
Only BPM
BPM with Pipe only
Simple but not realistic

## Slide 8
231 mm
110.5 mm
Dipole
Magnet
Dipole
Magnet
Pipe
Bellows
BPM
Bellows
BPM
Dipole
Magnet
Vacuum Duct
Realistic Model
BPM + Pipe + Bellows

## Slide 9
a. Mapping in CST Only with BPM
Similar condition with the Mapping setup
In the CST simulation, Ports are closed at the upstream and downstream ends. However, in Mapping, no dummy is used! Details will come later
b. Exaggerated condition!
Amplify Edge effect! Due to an abrupt change in the Beam Duct Diameter. No Linearities and a large offset are observed
c. Realistic Case as in RCS!
Kx/y =110 mm
Only ¼ region
is solved
a
b
c

## Slide 10
Fitting with 3rd Order Polynomial
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
5.45 | -5.42 | 161.30 | 161.73
CST Mapping: Only BPM
BPM’s Intrinsic transverse offset correction with a polynomial fit is about 5.5 mm.
Transverse offsets are nearly symmetric

## Slide 11
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
6.05 | 4.88 | 162.11 | 155.02
CST Mapping: BPM and Pipe (Exaggerated Condition)
Transverse offsets are NOT Symmetric!
Due to the upstream vacuum duct!!

## Slide 12
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
3.94 | -5.54 | 157.91 | 165.11
CST Mapping: BPM, Bellows and Pipe (Realistic Condition)
Transverse offsets are NOT Symmetric!
Due to the upstream vacuum duct!!

## Slide 13: What do we measure at RCS?
The mapping function is obtained from the wire mapping only with the BPM
In the RCS BPM upstream, the bellows and dipole magnet vacuum duct (sort of elliptical shape) is not considered!
  The edge effect in BPM due to the vacuum duct shape is not considered
  Results in huge errors?
 | C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
BPM | 5.45 | -5.42 | 161.30 | 161.73
BPM + Pipe | 6.05 | 4.88 | 162.11 | 155.02
BPM + Bellows + Pipe | 3.94 | -5.54 | 157.91 | 165.11
To estimate these errors in CST Simulation: Let’s consider the Xraw = Log[A/C] and Yraw=Log[B/D] from the case BPM + Bellows + Pipe, but the Fitting function from the BPM only!!

## Slide 14: Xraw and Yraw from the realistic case, but the Fitting function from the BPM only!! The error is 1.6 mm to 3.5 mm within the measurement range from 0 mm to ±55 mm
Error estimation

## Slide 15: Wire Mapping

## Slide 16: Wire Mapping System
Bellows
BPM
X/Y Mover
BPM and wire (Molybdenum: Ød=100 µm) aligned with accuracy of 100 µm
  Meaning: The mechanical center of BPM is known within 100 µm uncertainty
Transverse mover precision is 5 µm (Measured)
Transverse mapping range is ±60 mm
Tektronix signal generator (AFG31000)
  Vpp=5 V, Frequency = 1 MHz (Varied)
Bergoz-BB Electronics is used
  Dynamic range: -70 dBm to +5 dBm
  BW: 10 MHz
  ADC Bit: NI Digitizer 16-bit
A thermocouple is attached to the outer body of BPM to monitor temperature

## Slide 17
Step size: 1 mm
Frequency: 1 MHz
Repeat 17 Times
The precision of the mapping system? 300 µm
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
Wire Mapping: 17× Repetition

## Slide 18
Old Data with 10 mm step
New Data with 1 mm step
Wire Mapping: Frequency Dependence

## Slide 19
Wire Mapping: Frequency Dependence
T. Miura, PASJ2006, WP67
fRF varies from 1.02 to 2.44 MHz
Signal Generator Freq. varied to measure the BPM response
Strong frequency dependence is observed
Offset(y) varies from 5.5 to 10 mm in Freq. range 0.5 to 10 MHz
3rd order Polynomial Fit
High frequency results in reduced isolation between electrodes
Consequently, a bigger transverse offset
Asymmetric, transverse offset due to the bellows in the horizontal direction

## Slide 20
Wire Mapping: Temperature Logging during mapping
Spikes due to DAQ troubles?
Temperature variation within one measurement is 0.1 OC

## Slide 21: Source of Errors in Transverse Position Measurement by Diagonal Cut-Plane BPM
Asymmetric Transition between Vacuum Duct and BPM
  Edge Effect
Frequency De

## Slide 22: Back up

## Slide 25
0.5 MHz 1 MHz 1.5 MHz 2 MHz 2.5 MHz 3 MHz 4 MHz 5 MHz 6 MHz 7 MHz 8 MHz 9 MHz 10 MHz

## Slide 26
0.5 MHz 1 MHz 1.5 MHz 2 MHz 2.5 MHz 3 MHz 4 MHz 5 MHz 6 MHz 7 MHz 8 MHz 9 MHz 10 MHz