## Slide 1: Mechanical Design of CEPC Detector Magnet
Hou Zhilong

## Slide 2: Contents
Overview & Key Parameters
Magnet Assembly & Suspension
Thermal Management
Mechanical Analysis (FEA)
Challenges & Future plan

## Slide 3: Magnet Overview
Critical Parameters
Items | Parameters
Cryostat | 
Outer Diameter | 8470
Inner Diameter | 7070
Length | 9050
Coil | 
Outer Diameter | 7916
Inner Diameter | 7300
Length | 8150
Mass of coil | 185 t
Position of Magnet in detector
Cross section of Magnet

## Slide 4: Magnet Overview
Current leads channel
Gas-liquid separator
Cryostat
Thermal shield
Coil
Coil suspension
3D Model of Magnet
3D Cross Section of Magnet
Multi-Layer Insulation
  Out coil : 10~15 layers
  Out thermal shield：35~ 40 layers

## Slide 5: Coil Assembly & Suspension
Structure:
3-segment assembly (bolted to support cylinder), 10~15 layers MLI
Suspension:
  Axial: 8 tension rods/end
  Radial: 2 upwards + 4 diagonal rods/end
Rod Material: Titanium alloy
Total weight: ~185 t
Electromagnetic force resulting from coil offset
Helium tank
Axial rod
Helium pipe
Upwards rod
Helium tank
Radial rod
3D Model of Magnet Coil

## Slide 6: Coil Assembly & Suspension
Axial Rod
Upwards Rod
Radial Rod
Rod | Diameter（mm） | Length（mm） | Quantity
Axial | 30 | 4520 | 16
Radial | 32 | 1700 | 8
Upwards | 40 | 1650 | 4
Axial rods: Slender and long
Load cases:
Cold mass: 185 t
Lifting acceleration: 0.4g (X), 1.5g (Y), 0.6g (Z)
Impact forces: 740 kN, 2775 kN, 1110 kN
Electromagnetic forces (10mm offset: 391 kN; 15mm: 900 kN)
3D Model of Magnet Coil
Rod | Assembly | Cooling | Max Impact
Axial | 50 | 193 | 291
Upwards | 327 | 345 | 513
Radial (Up) | 120 | 224 | 314
Radial (Down) | 20 | 164 | 314
Tensile Stress of Rod
Parameters of Rod
MPa

## Slide 7: Coil Support Cylinder
Dimensions：Ø7820 mm,30 mm thick,2750 mm length
Material：5083 Aluminum alloy
Challenge: Large diameter + thin walls → Precision manufacturing

## Slide 8: Thermal Shield Design
Design: Inner/outer shields with serpentine cooling pipes
Suspension:
  6 diagonal rods/end
  Room temperature end offset towards the center
Rod Material: Titanium alloy
MLI:35~40 layers
Rod
Cooling Pipe
3D Model of Thermal Shield

## Slide 9: 2D Drawing of Magnet

## Slide 10: Cryostat Mechanical Analysis (FEA)
FEA Model
Material: Stainless steel
  E=190 GPa,
  µ=0.28,
  σ0.2 ≥205 MPa
Boundary conditions:
  Left end: 6 position fixed fully
  Right end: 6 position Horiz/vert fixed, axial free
Load Cases: Gravity, impact force caused by hoisting, vacuum, cooling, EM forces resulting from coil offset

## Slide 11: Cryostat Mechanical Analysis (FEA)
Loading
Axial load
Radial load

## Slide 12: Cryostat Mechanical Analysis (FEA)
Assembly
  Weight of cold mass
  Pre-stress of tie rod
Usummax=0.22 mm
Seqvmax= 51 MPa
Deformation
Von Mises Stress

## Slide 13: Cryostat Mechanical Analysis (FEA)
Axial accelerated motion
  Weight of cold mass
  Pre-stress of tie rod
  Axial impact force
Usummax=1.02 mm
Seqvmax= 130 MPa
Deformation
Von Mises Stress

## Slide 14: Cryostat Mechanical Analysis (FEA)
Vertical accelerated motion
  Weight of cold mass
  Pre-stress of tie rod
  Vertical impact force
Usummax=0.34 mm
Seqvmax= 80 MPa
Deformation
Von Mises Stress

## Slide 15: Cryostat Mechanical Analysis (FEA)
Radial accelerated motion
  Weight of cold mass
  Pre-stress of tie rod
  Radial impact force
Usummax=0.22 mm
Seqvmax= 51 MPa
Deformation
Von Mises Stress

## Slide 16: Cryostat Mechanical Analysis (FEA)
Vacuum
  Weight of cold mass
  Pre-stress of tie rod
  Air pressure
Usummax=0.76 mm
Seqvmax= 53 MPa
Deformation
Von Mises Stress

## Slide 17: Cryostat Mechanical Analysis (FEA)
Cooling
  Weight of cold mass
  Pre-stress of tie rod
  Air pressure
  Cooling shrink
Usummax=0.82 mm
Seqvmax= 94 MPa
Deformation
Von Mises Stress

## Slide 18: Cryostat Mechanical Analysis (FEA)
Excited , 15 mm misalignment in axial direction
  Weight of cold mass
  Pre-stress of tie rod
  Air pressure
  Cooling shrink
  EM force
Usummax=1.3 mm
Seqvmax= 134 MPa
Deformation
Von Mises Stress

## Slide 19: Key FEA Results
Condition | Max Deformation mm | Max Von Mises Stress Mpa
Assembly | 0.22 | 51
Axial accelerate | 1.02 | 130
Radial accelerate | 0.22 | 51
Vertical accelerate | 0.34 | 80
Vacuum | 0.76 | 52.6
Cooling | 0.82 | 93.7
Excited , 15 mm misalignment | 1.33 | 134
Deformation and Von Mises Stress

## Slide 20: Summary & Future Plan
Key Design Features
Modular Coil Assembly: 3-segment bolted structure, titanium alloy suspension system.
Critical Components: Precision-machined 5083 aluminum support cylinder.
Mechanical Performance
Load Cases: Gravity , impact force, vacuum, cooling, and electromagnetic forces (up to 900 kN at 15mm offset).
Results:
  Max stress: 134 MPa (during excitation with 15mm offset) – well below yield (205 MPa).
  Max deformation: 1.3 mm (within tolerance for detector alignment).
Key Challenges
Manufacturing: Thin-walled cylinder ,and slender suspension rods (L/D > 50).
Future plan
Design Refinement: Optimize rod parameter.
Prototyping: Validate machining processes for titanium rods and aluminum alloy cylinder.
Integration: Finalize cryostat welding and assembly protocols.
Conclusion
The design meets all mechanical requirements for CEPC detector operation. Focus areas include addressing manufacturing tolerances and stress concentrations, with prototyping to de-risk production.

## Slide 21
Thanks !