## Slide 1: Investigation of Diagonal Cut-Plane BPM Performance in the CSNS RCS
M. A. Rehman, Xu Zhihong, Qiu Ruiyang, Renjun Yang, Yang Rui, Weiwen Chen
2025-05-20
CSNS-BI

## Slide 2: Content
Reveal the Edge Effect in RCS BPM
  The EM field of the beam is distorted in the vicinity of the edges of the BPM [1]
    J-PARC MR also reported a similar issue [2, 3]
  Simulations
    CST simulations to reveal the edge effect
Calibration of the BPM with wire mapping
    Frequency Dependence
    Wire Mapping without Dummy pipe
    Wire Mapping with Dummy pipe
    Temperature influence on BPM
[1] J. H. Cupérus, NIM145 (1977) 233-243
[2] T. Toyama, IPAC10 MOPE012
[3] ] T. Toyama, OHO School 2009 (Japanese)

## Slide 3: Edge Effect in BPM
If the beam pipe and BPM cross sections/shapes are different
  The EM field of the beam is distorted in the vicinity of the edges of the BPM
  Design a BPM in such a way as to prevent the edge effect
      Guard rings in an upstream and downstream direction
      The longitudinal space is limited!
  Wire Mapping
      The wire is a constant-potential surface
      For the beam, the potential along its length depends on the charge of the beam and the cross-section of the beam pipe
      Underestimate the edge effect
  Mapping with Beam [only one at ANL ZGS (J. R. Simanton, 1969)]!
    Expensive and difficult to realize the installation
  Analytical formulae
    Less accurate, valid only for simple geometries
  FEM simulations (CST)
    Wake field solver line charge only few weeks!
    PIC solver beam transverse cross-section few months
J. H. Cupérus, NIM145 (1977) 233-243

## Slide 4: CST Simulation

## Slide 5: CST Simulations
Only BPM
  Bunch charge 1 µC, Bunch Length 7000 mm/1σ (24 ns/1σ ), β = 1
    ¼ Region, 0 to x/y:60 mm, with 10 mm step
BPM with Upstream Bending Magnet Pipe (Exaggerated Condition)
  Bunch charge 1 µC, Bunch Length 7000 mm, β = 1
    ¼ Region, 0 to x:60 mm, y: 40 mm, with 10 mm step
BPM with Upstream Bellows and Bending Magnet Pipe (Realistic)
  Bunch charge 1 µC, Bunch Length 7000 mm, β = 1
    ¼ Region, 0 to x:60 mm, y: 40 mm, with 10 mm step

## Slide 6
C11 C12 C13 C14
C21 C22 C23 C24
C31 C32 C33 C34
C41 C42 C43 C44
(
(
=
296.62 pF -8.56 pF -3.07 pF -1.40 pF
-8.56 pF 300.1 pF -7.50 pF -3.07 pF
-3.07 pF -7.50 pF 300.44 pF -8.54 pF
-1.40 pF -3.07 pF -8.54 pF 296.60 pF
220 mm
226 mm
CST Simulations: Capacitance and S-Parameters
Electrode Isolation
RCS BPM Model
Capacitance Matrix

## Slide 7
220 mm
226 mm
83.96 mm
24 mm
77 mm
135 mm
218 mm
3 mm
Only BPM
BPM with Pipe only
Simple but not realistic
CST Simulated Models

## Slide 8
231 mm
110.5 mm
Dipole
Magnet
Dipole
Magnet
Pipe
Bellows
BPM
Bellows
BPM
Dipole
Magnet
Vacuum Duct
Realistic Model
BPM + Pipe + Bellows

## Slide 9
a. Mapping in CST Only with BPM
  Similar condition with the Mapping setup
  In the CST simulation, Ports are closed at the upstream and downstream ends. However, in Mapping, no dummy is used! Details will come later
b. Exaggerated condition!
  Amplify Edge effect! Due to an abrupt change in the Beam Duct Diameter. No Linearities and a large offset are observed
c. Realistic Case as in RCS!
a
b
c
Each case results in a different offset!
The mapping function obtained from each case will be Different!
CST Mapping

## Slide 10
Fitting with 3rd Order Polynomial
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
5.45 | -5.42 | 161.30 | 161.73
CST Mapping: Only BPM
BPM’s Intrinsic transverse offset correction with a polynomial fit is about 5.5 mm.
Transverse offsets are nearly symmetric

## Slide 11
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
6.05 | 4.88 | 162.11 | 155.02
CST Mapping: BPM and Pipe (Exaggerated Condition)
Transverse offsets are NOT Symmetric!
Due to the upstream vacuum duct!!

## Slide 12
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
3.94 | -5.54 | 157.91 | 165.11
CST Mapping: BPM, Bellows and Pipe (Realistic Condition)
Transverse offsets are NOT Symmetric!
Due to the upstream vacuum duct!!

## Slide 13: What do we measure at RCS?
The mapping function obtained from the wire mapping ignores the BM pipe!
  The edge effect due to the vacuum duct shape is ignored
  Results in huge errors?
To estimate these errors in the CST Simulation
  Let’s consider the Xraw = Log[A/C] and Yraw = Log[B/D] from the case BPM + Bellows + Pipe
  But the Fitting function from the BPM only!!
 | C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
BPM | 5.45 | -5.42 | 161.30 | 161.73
BPM + Pipe | 6.05 | 4.88 | 162.11 | 155.02
BPM + Bellows + Pipe | 3.94 | -5.54 | 157.91 | 165.11

## Slide 14
Error estimation: CST Simulation
Xraw and Yraw from the realistic case
but the Fitting function from the BPM only!!
The error is 1.6 mm to 3.5 mm within the measurement range from 0 mm to ±55 mm

## Slide 15: Wire Mapping

## Slide 16: Wire Mapping System
Bellows
BPM
X/Y Mover
BPM and wire (Molybdenum: Ød=100 µm) aligned with accuracy of 100 µm
  Meaning: The mechanical center of BPM is known within 100 µm uncertainty
Transverse mover precision is 5 µm (Measured)
Transverse mapping range is ±60 mm
Tektronix signal generator (AFG31000)
  Vpp=5 V, Frequency = 1 MHz (Varied)
Bergoz-BB Electronics is used
  Dynamic range: -70 dBm to +5 dBm
  BW: 10 MHz
  ADC Bit: NI Digitizer 16-bit
A thermocouple is attached to the outer body of BPM to monitor temperature

## Slide 17
Step size: 1 mm
Frequency: 1 MHz
Repeat 17 Times
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
Wire Mapping: 17× Repetition

## Slide 18
Wire Mapping: Frequency Dependence

## Slide 19
Temperature Logging during mapping
Temperature variation within one measurement is 0.1 OC

## Slide 20
Wire Mapping: Frequency Dependence
T. Miura, PASJ2006, WP67
fRF varies from 1.02 to 2.44 MHz
Signal Generator Freq. varied to measure the BPM response
Strong frequency dependence is observed
Offset(y) varies from 5.5 to 10 mm in Freq. range 0.5 to 10 MHz
3rd order Polynomial Fit
High frequency results in reduced isolation between electrodes
Consequently, a bigger transverse offset
Asymmetric, transverse offset due to the bellows in the horizontal direction

## Slide 21
Room Temperature Variation influence on BPM!
Room Temperature
23OC
Room Temperature
27OC
One measurement
26.33±0.11 OC
26.59±0.02 OC
26.71±0.02 OC
26.75±0.004 OC
22.79±0.04 OC
22.84±0.02 OC

## Slide 22
Room Temperature Variation Influence on BPM!
A Very Drastic Effect has been observed in the Transverse offset (especially in the low-frequency region) due to
  Influence of temperature rise on BPM electrodes and body
  Influence of temperature rise on Electronics

## Slide 23
30.26±0.07 OC
Beam ON
Temperature sensor
BPM
Bending Magnet
Temperature sensor Readout
The temperature sensor is about 0.5 meters away from the BPM
Near BPM Temperature in Tunnel

## Slide 24
Dummy-pipe (Edge Effect) influence on BPM!
Dummy Pipe
BPM
Dummy Pipe

## Slide 25
Dummy-pipe (Edge Effect) influence on BPM!
A Very Drastic Effect has been observed in the Transverse

## Slide 26: Summary
Frq [MHz]
Cx,00
23OC
26OC
23OC W/ Dummy Pipe
Cy,00
Frq [MHz]
23OC
26OC
23OC W/ Dummy Pipe
Very strong dependence of the calibration constant
  Frequency
  Temperature

## Slide 27: Back up