## Slide 1: Bergoz-BB tests
Rehman, Weiwen, Qiu, Xu, Renjun, YangRui, ChenMing

## Slide 2: Bergoz BB electronics Check
Devices
  Input Generator: RIGOL DG922Pro (200 MHz, 1.25 GS/s)
    Input signal: Sine wave
  Keysight oscilloscope DSOX3054A, 500 MHz, 4 Gsa/s
    Averaging: 64
  RF Splitter 4 Way Minicircuit DC-8.4 GHz, ZFRSC-4-842-S+
  Attenuator: Minicircuit, VAT2+,3+,10+,20+
Measurement
  The RF splitter and attenuator frequency dependence were checked
    Independent of frequency (checked range: 0.5 MHz to 10 MHz)
  Bergoz BB electronics
    Linearity check (vary the attenuator at fixed frequency) @ 27OC and 23OC
      No significant dependence of linearity on temperature was observed
    Frequency dependence (Vary the frequency at fixed attenuator )
      No frequency dependence of the Bergoz-BB module was observed
      No significant dependence of frequency on temperature was observed

## Slide 3
Attenuator Test

## Slide 4: RF Splitter Test

## Slide 5: Laboratory Temperature
The Set Laboratory Temperature was 27°C
Measured Temperature near BB Module was 30°C
The Set Laboratory Temperature was 23°C
Measured Temperature near the BB Module was 27°C

## Slide 6
Bergoz-BB Frequency dependence Check
Bergoz-BB have a BW of 10 MHz
  So, it should be independent of the signal up to 10 MHz
Data was obtained
  Vpp: 5 mm, Frequency 0.5 MHz to 10 MHz @ 27 °C
    No dependence on Frequency
  Vpp: 100 mm, Frequency 0.5 MHz to 10 MHz @ 27OC and 23OC, - 4 dB att.
    No dependence on Frequency or Temperature
  Vpp: 100 mm, Frequency 0.5 MHz to 10 MHz @ 27OC and 23OC, +4 dB att.
    No dependence on Frequency or Temperature

## Slide 7
Bergoz-BB Linearity Check
Bergoz-BB possesses a dynamic range of dBm
    The difference between the two different temperatures is less than 2%
    No significant dependence on temperature

## Slide 8
Structural Deformation
under vacuum
The electrodes are body of the BPM is only 1.5 mm thin (made of Ti-alloy)
Rib structure is used to preserve mechanical stability
Stress analysis has been carried out in CST simulation to estimate the mechanical deformation under vacuum condition
The deformation under vacuum condition is only 15.6 µm!
Negligible!

## Slide 9
Mapping with Dummy Pipe and Higher Temperature