## Slide 1: CSNS RCS Diagonal Cut-Plane BPM: Accuracy vs. Calibration
M. A. Rehman
On behalf of the CSNS Beam Instrumentation group
Many thanks to CSNS Alignment & Survey, Mechanical, PS group
2025-05-29

## Slide 2: Content
Reveal the Edge Effect in RCS BPM
  The EM field of the beam is distorted in the vicinity of the edges of the BPM [1]
    J-PARC MR also reported a similar issue [2, 3]
  Simulations
    CST simulations to reveal the edge effect
Calibration of the BPM with wire mapping
    Frequency Dependence
    Wire Mapping without Dummy pipe
    Wire Mapping with Dummy pipe
    Temperature influence on BPM
[1] J. H. Cupérus, NIM145 (1977) 233-243
[2] T. Toyama, IPAC10 MOPE012
[3] T. Toyama, OHO School 2009 (Japanese)
[4] K. Satou, OHO School 2020 (Japanese)

## Slide 3: CST Simulations

## Slide 4: CSNS RCS BPM Structure
A
C
B
D
32 Diagonal cut BPMs are instrumented in RCS
Several of them are placed inside steering magnet (limited space)
  Rib structure on body to prevent eddy current
BPMs are 45O rotated for convenient access to attach cables
Corrector
Magnet
BPM
Diagonal Cut BPM design by
Prof. Taoguang Xu

## Slide 5: CST Simulations
Only BPM
  Bunch charge 1 µC, Bunch Length 7000 mm/1σ (24 ns/1σ ), β = 1
    ¼ Region, 0 to x/y:60 mm, with 10 mm step
BPM with Upstream Bending Magnet Pipe (Exaggerated Condition)
  Bunch charge 1 µC, Bunch Length 7000 mm, β = 1
    ¼ Region, 0 to x:60 mm, y: 40 mm, with 10 mm step
BPM with Upstream Bellows and Bending Magnet Pipe (Realistic)
  Bunch charge 1 µC, Bunch Length 7000 mm, β = 1
    ¼ Region, 0 to x:60 mm, y: 40 mm, with 10 mm step

## Slide 6
218 mm
Scenario 1: only BPM
CST Simulated Models: Three Scenarios
Scenario 2: BPM+short pipe
Scenario 3: BPM+realistic pipe
(R1BPM04 case)

## Slide 7
 | C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
Senario1 | 5.45 | -5.42 | 161.30 | 161.73
Senario2 | 6.05 | 4.88 | 162.11 | 155.02
Senario3 | 3.94 | -5.54 | 157.91 | 165.11
Evaluation of position map in simulation
To find offset and calibration constant polynomial fit
3rd Order Polynomial was used
The mapping function obtained from the wire mapping ignores the BM pipe!
To estimate these errors in the CST Simulation
  Let’s consider the Xraw = Log[A/C] and Yraw = Log[B/D] from the case BPM + Bellows + Pipe
  But the Fitting function from the BPM only!!

## Slide 8
Offset estimation in simulation
Xraw and Yraw from the realistic case
but the Fitting function from the BPM only!!
The accuracy error is ? mm to?mm within the measurement range from 0 mm to ±10 mm
R.YANG: Offset in hor. and ver. (NOT 45 deg tilt!)

## Slide 9: BPM calibration in lab.
Many thanks to the Alignment and Survey Group (J. Liang, T. Wang)

## Slide 10: Calibration system in D1
Bellows
BPM
X/Y Mover
BPM and wire (Molybdenum: Ød=100 µm) aligned with accuracy of 100 µm (Measured)
Transverse mover precision is 5 µm (Measured)
Transverse mapping range is ±60 mm
Tektronix signal generator (AFG31000)
  Vpp=5 V, Frequency = 1-10 MHz (Varied)
Bergoz-BB Electronics
  Dynamic range: -70 dBm to +5 dBm
  BW: 10 MHz
  ADC Bit: NI Digitizer 16-bit
A thermocouple to monitor temperature

## Slide 11: Survey of the calibration system
1.First Measure Coordinate of BPM
  Coordinate Measuring Machine（CMM）
  Brand: HEXAGON
  Model: PILOT 15.30.12
BPM calibration 01 |  |  |  |  | BPM calibration 02 |  |  | 
point name | X/mm | Y/mm | Z/mm |  | point name | X/mm | Y/mm | Z/mm
F1 | -88.214 | 185.064 | 244.628 |  | F1 | -88.218 | 185.070 | 244.626
F2 | -88.026 | 190.152 | 51.868 |  | F2 | -88.014 | 190.145 | 51.897
F3 | 87.936 | 189.844 | 51.456 |  | F3 | 87.993 | 189.804 | 51.465
F4 | 87.988 | 185.130 | 244.477 |  | F4 | 88.045 | 185.090 | 244.490
BPM calibration 03 |  |  |  |  | BPM calibration 04 |  |  | 
point name | X/mm | Y/mm | Z/mm |  | point name | X/mm | Y/mm | Z/mm
F1 | -88.211 | 185.074 | 244.630 |  | F1 | -88.207 | 185.077 | 244.628
F2 | -88.044 | 190.170 | 51.882 |  | F2 | -88.014 | 190.154 | 51.866
F3 | 87.944 | 189.838 | 51.457 |  | F3 | 87.948 | 189.838 | 51.459
F4 | 87.993 | 185.126 | 244.479 |  | F4 | 87.995 | 185.123 | 244.475
error | BPM calibration 01-BPM calibration 02 |  |  | 
 | X/mm | Y/mm | Z/mm | 
F1 error | -0.004 | 0.005 | -0.003 | 
F2 error | 0.012 | -0.007 | 0.030 | 
F3 error | 0.057 | -0.040 | 0.010 | 
F4 error | 0.057 | -0.039 | 0.013 | 
error | BPM calibration 04-BPM calibration 03 |  |  | 
 | X/mm | Y/mm | Z/mm | 
F1 error | 0.003 | 0.003 | -0.002 | 
F2 error | 0.030 | -0.016 | -0.015 | 
F3 error | 0.004 | -0.001 | 0.002 | 
F4 error | 0.002 | -0.002 | -0.004 | 
error | BPM calibration 04-BPM calibration 01 |  |  | 
 | X/mm | Y/mm | Z/mm | 
F1 error | 0.007 | 0.013 | -0.001 | 
F2 error | 0.012 | 0.002 | -0.001 | 
F3 error | 0.013 | -0.006 | 0.003 | 
F4 error | 0.007 | -0.006 | -0.002 | 
R.YANG: Conclusion of this page

## Slide 12: Survey of the calibration system
Chart: X
Chart: Y
X axis | 
error(forward-goning) | 
max | 0.03mm
min | -0.012mm
RMS | 0.006mm
error(reverse-goning) | 
max | 0.008mm
min | -0.042mm
RMS | 0.005mm
Y axis | 
error(forward-goning) | 
max | 0.022mm
min | -0.045mm
RMS | 0.017mm
error(reverse-goning) | 
max | 0.025mm
min | -0.065mm
RMS | 0.016mm
Laser Scanner
Brand: FARO
Model: S730
R.YANG: Conclusion of this page

## Slide 13: Survey of the calibration system
Laser Scanner
Brand: FARO
Model: S730
Error of BPM Center and Wire Position |  |  |  |  |  |  | 
Test 01 | X(mm) | Y(mm) |  | Test 02 | X(mm) | Y(mm) | 
Top Flange center of BPM-Wire | 0.018 | 0.038 |  | Top Flange certer of BPM-Wire | -0.079 | 0.02 | 
Bottom Flange center of BPM-Wire | -0.005 | 0.014 |  | Bottom Flange certer of BPM-Wire | -0.052 | 0.011 | 
Test 03 | X(mm) | Y(mm) |  |  |  |  | 
Top Flange center of BPM-Wire | -0.066 | 0.072 |  |  |  |  | 
Bottom Flange center of BPM-Wire | -0.088 | -0.018 |  |  |  |  | 
R.YANG: Conclusion of this page

## Slide 14: Test of the BPM electronics

## Slide 15
Bergoz-BB Frequency dependence Check
Bergoz-BB have a BW of 10 MHz
  So, it should be independent of the signal up to 10 MHz
Data was obtained
  Vpp=5 mV, Frequency 0.5 MHz to 10 MHz @ 27 °C
  Vpp=100 mV, Frequency 0.5 MHz to 10 MHz @ 27OC and 23OC, Log(Va/Vc) = -4 dB
  Vpp=100 mV, Frequency 0.5 MHz to 10 MHz @ 27OC and 23OC, Log(Va/Vc) = +4 dB
Frq [MHz]
Xout@27O
Xout@23O
The room temperature was adjusted with the air conditioner.
R.Yang: 0dB, -4 dB, +4dB are misleading
Abs. error to position.

## Slide 16
Bergoz-BB Linearity Check
Att. [dB]
Xout@27O
Xout@23O
Yout@27O
Yout@23O
R.YANG: Conclusion of this page
Log(Va/Vc) or Log(Vb/Vd)

## Slide 17: Re-calibration of the BPM

## Slide 18
Step size: 1 mm
Frequency: 1 MHz
Repeat 17 Times
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
C00,x [mm] | C00,y [mm] | C10,x [mm] | C01,y [mm]
Repeatability of the calibration system
Conclusion:
The random error induced uncertainty of the mapping coef. is <3.2%.

## Slide 19
Calibration system: more details
1, dummy pipe (old) were added to the upper and bottom sides
2, Room temperature was change from 23 oC to 27 oC (maximum!)
Input to Electronics: 1~2 mV

## Slide 20
Wire Mapping: Frequency Dependence
Strong frequency dependence is observed
Offset(y) varies from 5.5 to 10 mm in Freq. range 0.5 to 10 MHz
The offset varies largely with and without dummy pipe
The offset variation with temperature change (w/o dummy pipe) is large. Offset variation with temperature change (w/ dummy pipe) is small
* w/ 3rd order Polynomial Fit

## Slide 21: Summary
CST simulations has been carried out to estimate the Edge Effect
  The error is ?mm to ? mm within the measurement range from 0 mm to ±55 mm
The wire mapping of spare BPM has been carried out
  Wire mapping without dummy pipe
    Room Temperature 23O
      Frequency variation 0.5 MHz to 10 MHz
    Room Temperature 27O
      Frequency variation 0.5 MHz to 2 MHz
  Wire mapping with dummy pipe
    Room Temperature 23O
      Frequency variation 0.5 MHz to 5 MHz
    Room Temperature 27O
      Frequency variation 0.5 MHz to 5 MHz
Known
Only Sim.
Exp. In future
Mapping
Exp.
Mapping
Exp.
Probably Error w/ and w/o Dummy Pipe: About 2.5 mm/3.7 results in additional offset @ r=0 mm

## Slide 22
30.26±0.07 OC
Beam ON
sensor
BPM
Bending Magnet
Temperature sensor Readout
Real BPM Temperature in Tunnel

## Slide 23
Highlights from 29/05/2025 RCS BPM meeting
A remote discussion meeting with T. Toyama and K. Satou of J-PARC/KEK (29/05/2025)
Some highlights of the discussion
  Comment 1#: CST simulation with open boundy might include additional undesired artifical effects/errors. Therefore, simulations with dummy pipe is recommanded. The appropriate length of the dummy pipe should be checked via simulation and comparison of the results. => A length of the dummy pipe being >3*BPM diameter has been recommanded by senior master; but, Toyama-san think L>1*diam. Might also work. It should be confirmed through comparisons.
  Q6: Did you check the capacitance matrix terms in lab.? A6: Not yet, but it could be done. Recommendation: use a two-port VNA and a 50 Ohm terminator; The lower frequency range is more important, i.e., determine the C_ij value in a lower freq. (<1 MHz at JPARC , Toyama-san)
  Q9: which dummy pipe length is favoured for the calibration system in lab.? (R. Yang) A9: Same troubles in J-PARC’s BPM calibration (varying BPM apertures). No clear norm. The 3*Diam. is assessed by senior master through static calculation (formula), and it is very rough. (Toyama-san) => We should use a longer one

## Slide 24
Highlights from 29/05/2025 RCS BPM meeting
Some highlights of the discussion
  Q7: Mapping with VNA at JPARC. A7: we use a 5-port VNA Recommendation:
    Generator – Transformer at wire input – Wire: the transformer should be optimized to minimize the reflection; in the case of mis-matched impedance, the reflection RF signal will cause extra EM field distortion inside the BPM body, it does deform the waveform
    Consider to use a RF power amplifier
    might be 200-300 Ohm characteristics impedance of the (wire-BPM, the calibration system)
    In Toyama-san’s mapping work, some problems were observed in the early phase, and the reason was found to be the mis-match between wire and coax. cable

## Slide 25: Future Works
Wire Mapping with realistic pipe (preparing)
Upgrade of the BPM calibration system
  A longer dummy pipe (defined in simu.)
  Controlled room temperature (up to 30 oC)
  Frequency of interest: 0-10 MHz
  Fasten the calibration process w/ VNA
Try to understand the real cause of strong frequency dependence
  Is it due to cross coupling among electrodes?
  Or Transfer Impedance?

## Slide 26: Back up

## Slide 27: Analysis Method
In case of non rotated electrodes
  X = cxLog[A/C], Y = cyLog[B/D]
  X and Yr come directly come from Bergoz-BB electronics,
  cx and cy are calibration constant
Since the electrodes are 45O rotated
  X = cx [Log(A/C) . cos(α) – Log(B/D) . sin(α)], Y = cy [Log(A/C) . sin(α) + Log(B/D) . cos(α)]
  sin(α)= cos(α)=1/sqrt(2)
In our case 3rd order polynomial is applied