## Slide 1: ITk strip hybrid and module production
Weiguo Lu（陆卫国）
On behalf of the ATLAS ITk China group
MOST ATLAS Detector Upgrade Project
IHEP, 5 Dec. 2025

## Slide 2: Outline
Introduction
Production Flow
  Assembly
  Qualification Control
  Wire-bonding
  Electrical test
Progress and Status
  Progress of the whole Collaboration
  Local status and risks at IHEP
Summary

## Slide 3: Module QC
Modules transitioned through a variety of stages and QC tests to ensure a working module
  Modules Visually inspected at each stage
    Visual inspection tests uploaded only if there is a problem
  Modules can move to the failed stage at any point
Descoping
  Tests/procedure not providing useful data are removed from the QC chain
  Glue weight , Pre-TC Bow…
Grading
  Some module metrology are out of the spec but believed installable, new Grade B
Limbo stage for EBD modules with hopeful recovery

## Slide 4: bPOL issue
Two kinds of failure modes
  A “gradual” failure
  A “sudden” failure
Tests and simulations to understand the reason
Tests to determine the radiation level reachable
  Verify if can be used reliably below ~5Mrad, 10Mrad…
Corrective solutions
  Modify the powerboard
  Modify the chip bPOL12V itself, new fabrication
Keep the current rate and no ramping up for ITk production

## Slide 5: Module Production @ IHEP
Assembly Progress
  Hybrids — 65
  Pre-production Modules (PPC + Pre-series) — 5 + 15
  Production Modules — 4 all passed
PPC not included in the plot
Short halt due to lack of glue from August to November

## Slide 6: Qualification Control
Hybrid
module

## Slide 7: Production Tools
A couple of new tools made in China
  1. ASIC tray
  2. Hybrid assembly Jig
  3. Pickup tool
  4. Module assembly Jig
  5. Module bonding Jig
  6. Hybrid panel Jig
  7. Module test Jig
Performed the quality check for all tools
  Training with dummy parts
Glass dummy
Mechanical dummy

## Slide 8: Coldjig improvement
The noise was higher at some chucks during the Pre-series production
The test results were analyzed to investigate the reason
A lot of efforts contributed to the setup improvement
  Re-assemble of the chucks
  Re-organize the layout of the system
  Better grounding and shielding
The MTC is now working as expected
  The noise of all chucks below ~1000e-
PS-4 away chuck4 227

## Slide 9: Potential Difficulties-assemble
Flex Thickness Variability Cripples Assembly Efficiency
  Root Cause：Attaching interposers to hybrids & PBs induces major batch-to-batch flex thickness variation
  Compounding Problem：Thickness is also non-uniform within each flex piece
  Operational Consequence：Requires tooling recalibration for every single batch
  Final Impact：Severely reduces assembly efficiency and introduces unavoidable process errors
Current Production Bottleneck & Proposed Solution
  Capacity Limit : The current lab vacuum system can only support the assembly of 3 modules per week.
  Throughput Goal: To increase the production rate, we need to assemble 2 modules simultaneously.
  Key Requirement: This parallel processing necessitates an additional, dedicated vacuum system to sustain the operation.

## Slide 10: Potential Difficulties-test
Deionization
  The static charge accumulated at the suction cup position
  Influence the noise measurement of the module
  A lot of time needed to remove the SCN
Degradation of Coldjig
  Aging of key parts of the system
    e.g. the humidity sensor
    The compressed air machine maintenance

## Slide 11: Summary
ITk strip barrel module production is challenging
IHEP passed site qualification and enter Production phase
Efforts made on QC including SOP development
Produced several modules which makes production flow much smoother
Difficulties ahead and mitigation methods on way
Thanks

## Slide 12: Back up

## Slide 13: plan

## Slide 14: Outline
Introduction
Production Flow
  Assembly
  Qualification Control
  Wire-bonding
  Electrical test
Progress and Status
Summary

## Slide 15: ITk Detector
HL-LHC upgrade
All silicon tracker
Barrel-Stave-SS/LS module
Endcap-Petal-R0/R1/R2/R3/R4/R5 module
Barrel
Endcap
SS module

## Slide 16: Barrel Module
Sensor
  Size : 97mm✖97mm
  Strip length LS: 48.2 mm / SS: 24.1 mm
  Pitch :75.5um
Hybrid
  ABCStar: ATLAS front-end Binary read-out Chip
  HCCStar: Hybrid Controller Chip
Powerboard
  AMAC: monitor and control
  bPOL12V: low-voltage regulation
  linPOL12V: both voltages for AMAC
  HVmux: high-voltage switching
Glue for mechanical fixing
Wire-bonding for electrical connection

## Slide 17: Outline
Introduction
Production Flow
  Assembly
  Qualification Control
  Wire-bonding
  Electrical test
Progress and Status
Summary

## Slide 18: Production Flow
Module Production
  Hybrid Assembly
  Module Assembly
  Quality Control (QC)
Site Qualifications
  IHEP have been qualified for all 29 QC steps in barrel module production
  Finished the Pre-Production and started production
Gluing area
Bonding area
E-test area

## Slide 19: Hybrid Assembly
Pick up ASICs and
alignment
Apply UV glue on hybrid and ASIC attachment
using dot pattern
UV glue curing
Wire bonding
Procedures
  ASICs attachment
    ASICs adhesive: UV-curable glue (Loctite 3525)
    Glue coverage controlled by weight and thickness
  Wire-bonding
    Hesse Bondjet BJ820
    25µm diameter aluminum wires
    Electrical connection between ASICs and hybrid flex (as well as hybrid and test panel)

## Slide 20: Module Assembly
Dispense glue on stencil for
hybrid
Dispense glue on stencil for
PB
Glue curing for hybrid and PB
Wire bonding
Procedures
  Hybrid / Powerboard attachment
    Hybrid and PB adhesive: Epoxy (True Blue)
  Wire-bonding
    Hesse Bondjet BJ820
    256 Al wires in 4 rows per ABCStar chip at Front-end
    Hybrid to powerboard wire-bonding
    Module to test frame wire-bonding

## Slide 21: Qualification Control
Glue Weight
  ASICs-to-hybrid glue dispensing by a 3-Axis E Series Automated Fluid Dispensing Robot
  Weighting with a digital scale
  43.80±2.62 mg for ASIC adhesive weight
  Calibration with dispensing time and air pressure
A Norson dispenser + JANOME robot
A Mettler digital scale

## Slide 22: Qualification Control
Metrology
  Geometric positioning of characteristics
    ASIC glue height: 120 +40/-60μm
    ASIC xy-placement: within +/-200μm
    ASIC flatness: Tilt ≤ 0.025
    Hybrid package thickness: 0.84±0.04mm
    Module bow between -50μm and +150μm
ASIC relative position
Glue height
Total package height
ASIC tilt
Glue height
Maximum height
Relative position

## Slide 23: Qualification Control
Visual Inspection
  Check hybrid flex
    Defects
  Check ASICs
    Integrity of surface and edges
    Cleanness on bonding pads
  Check Sensor
    Any scratches, marks and debris
    Integrity of edge
    broken sensor
  Check bond-wires
FE wire-bondings
ASICs
sensor
Suction cup marks
Long scratch
A Keyence digital microscope

## Slide 24: Wire Bonding
Module wirebonding
  Front end – four columns of pads with 75.5um pitch
  Yield of the electronics channel >99%
Bonder Hesse BJ 820
Nordson Dage 4000Plus
Pull test every month to optimize the bonding parameters
  Pull force > 8g
  Peel break (4/5) < 10%
  Variation of different surfaces

## Slide 25: Electrical Test
Reception of sensor and powerboard
  Required to ensure the quality
  Powerboard
    alive and functionality test
  Sensor
    IV test based on single module test system

## Slide 26: Electrical Test
Hybrid Burn-in
  The hybrids are required to be burned in before using for module assembly
  100（20） hours running to check hybrid operation stability at ~40 ℃
  Voltage and current of AMAC, temperature of AMAC and hybrid are monitored by InfluxDB/Grafana
Hybrid burn-in performs cycles of ITSDAQ Full Test
  Pedestal Trim
  Strobe Delay
  Three Point Gain
  Response Curve
  Noise Occupancy

## Slide 27: Electrical Test
Module Thermal Cycling
  Prior to Stave/Petal assembly and evaluate each module to determine if it should go forward to Stave/Petal assembly
  Temp/Humidity/LV/HV/interlock/data readout
  Each module undergoes electrical and thermal quality control (QC) testing
    Thermal-cycled 10 times (-35℃ ~ +20℃, +20℃ ~ -35℃)
    AMAC IV scan (0, -550V)
    ITSDAQ Full Test carried at -350V performed at -35℃ and +20℃ each cycle
    2 hours HV stability measurement and Open Channel Search test is performed at -10V

## Slide 28: IHEP SOP
IHEP ATLAS-ITK Standard Operating Procedure
  Environment real time monitoring of cleanroom
    Temperature and humidity
    Dust Particle
  Local production operation guidance
  Local production statistics and report
  Production database interaction