## Slide 1: A FPGA-based DAQ system for RPC cosmic ray test
Jiaxuan Li
On behalf of USTC RPC lab
Dec 5th, 2025

## Slide 2: Outline
Introduction
FPGA-based DAQ system design
Performance of the TDC with the signal generator
DAQ test with RPC detector
Summary

## Slide 3: Introduction

## Slide 4: Introduction
Phase II upgrade: next-generation Resistive Plate Chambers
  < 300 ps time resolution
  1 cm spatial resolution
  > 95% detection efficiency
Cosmic ray test in RPC production
  monitor singlets
  test singlets
Advantages
  monitor singlets are very close together → a better triggering rate (try to accept all hit angles)
  can operate the test singlets easliy
DAQ design goals
  144 channels per board
  TDC function with 100 ps time resolution
  three trigger modes: external/self/master-slave

## Slide 5: FPGA-based DAQ system design
2025 CLHCP

## Slide 6: Overall signals system
DAQ:
Located or directly on the chamber or maximum at 3.5m distance, connected with 18 FEBs.
FEBs eta1 layer 0
Signals ‘’spyder’’ board
FEBs eta1 layer 1
FEBs eta1 layer 2
20 pins cables
Differential BC_clock cable
68 pins cables
Signals ‘’spyder’’ board
The Signals ‘’spyder’’ board is an adaptor board with 4 connectors.
Located directly within the service boxes of the detector
It couples in a single 68 pins cable the signals of 3 FEBs of a readout panel column
Front-end board
Backup solution: without TDC
Amplifier and discriminator
20 pins connect with spyder board
Low voltage board
Power supply for FE
Vth / Vdigital / Vamp / Vpol
Low voltage board
PC software
TDC data transmission
Online monitor

## Slide 7: DAQ hardware design
Receive data from FE board
18 FE boards = 144 channels in total.
Integrates 144 channels polarization circuitry. Capable of receiving signals from baseline or backup FE
Xilinx KINTEX UltraScale: XCKU060
Central processing and control Manchester decoder or TDC
CPLD:XC95144XL
Processes 18 fast_or signals from FE board.
JTAG
SFP
UART
Ethernet
Clock buffer
40 MHz LVDS clock output from FPGA.
Split into 2 channels,
then further distributed into 18 channels.
Other
Temperature sensor, SMA, buttons, fan, cascading signals between two boards, and etc.

## Slide 8: TDC principle diagram
Input sample part
  Phase shift sampling clock: 500MHz
  8 bit thermometer code
Analysis part
  8 bit → 3 bit fine count
  Analyze edges
Trigger logic
  External/self/master-slave
Ethernet communication part
  Transfer the data to the total FIFO
  Gigabit Ethernet

## Slide 9: Software interface
🟩 Binary → ROOT conversion
One-click generation of ROOT files
Supports large continuous runs
Used for offline analysis
🟧 Online monitor (in progress)
Time-window configuration
Live hit map / efficiency / pulse-width display
Event counter and efficiency display
🟪 Status
✔ Data-taking functions fully verified
✔ Binary-to-ROOT conversion operational
⏳ Online monitor being tested
🟦 DAQ data-taking functions
Ethernet-based raw data reception
Stable binary streaming from 4 TDC modules
File saving with size/quantity display

## Slide 10: Performance of the TDC with the signal generator
2025 CLHCP

## Slide 11: Performance: DNL & INL
Test setup
  Two pulse signals as trigger and stop
  Standard deviation is the time resolution
Test results: DNL & INL
General Pulse Width: 2 – 21 ns Trigger delay: 400 ns
.
Max DNL = + 0.18 LSB
Max INL = +0.23 LSB
Max = 0.4 LSB
Result: The impact of bin width variation on resolution is non-dominant compared to quantization error, indicating that the linearity performance of the TDC is well within acceptable limits.

## Slide 12: Performance: time resolution
Test results: time resolution
(c)
(d)
Each channel’s mean value shows uniformity after applying the correction.
All channels achieved a time resolution of 89 ps.
FPGA-based TDC attains sub-100 ps time resolution and 2 ns dead time over 144 channels.

## Slide 13: DAQ test with RPC detector
2025 CLHCP

## Slide 14: Performance with RPC detector: setup
Triplet BIS RPC
2 scintillators
6 FEEs
Low voltage board
Spyder board
2 DAQ boards
2 scintillators
HV control
Trigger System: two scintillators operate in coincidence to generate the external trigger signal.

## Slide 15: Performance with RPC detector: efficiency curve
90
80
60
50
40
30
•
First efficiency curve measured
using the USTC DAQ entirely
Without the Faraday cage
20
•
Vth = 1,8V Vamp = 1,2V
10
0
4500
4700
4900
5100
5300
5500
5700
5900
6100
HV (V)
Efficiency (%)
The measurements were taken after improving the Faraday cage, with updated FE parameters.
Efficiency taken with the Dis_OR and all the cable plugged to the DAQ in operation and grounded together
Vth = 2.1V Vamp = 1.45V
Efficiency measured in the same condition but using the USTC DAQ
Vth = 2V Vamp = 1.4V
Efficiency measured in the same condition but using the USTC DAQ
Vth = 1.9V Vamp = 1.4V
70
100
Notes: Without the Faraday cage, the noise level is high, requiring a higher threshold and resulting in reduced efficiency. After adding the Faraday cage, the noise is significantly suppressed, allowing the use of a lower threshold. The measured efficiency matches the observations from the oscilloscope and is consistent with expectations.

## Slide 16: Performance with RPC detector: results
Hit map matches the setup, with scintillators placed at the indicated positions.
Pulse widths are 5–10 ns and the cluster-size mean is about 1.3.

## Slide 17: Summary
2025 CLHCP

## Slide 18: Summary
DAQ module: TDC key functions have been verified.
Readout system: full-chain test completed.
Software: Data acquisition works stably; the online monitor still needs full validation.
Next step:
  Verify self-trigger modes.
  Complete event matching and perform position reconstruction analysis.
  Finalize the online monitor features.
Thank you!

## Slide 19: Back up
2025 CLHCP

## Slide 20: Architecture of data sampling
A pulse signal following fan-out is captured simultaneously by eight flip-flops driven by eight uniform phase-shifted clocks.
Flip-flops | First Level Fan-out Delay (ps) | Second Level Fan-out Delay (ps) | Third Level Fan-out Delay (ps) | Total Delay (ps)
1 | 305 | 239 | 370 | 914
2 | 305 | 239 | 363 | 907
3 | 305 | 140 | 403 | 848
4 | 305 | 140 | 365 | 810
5 | 268 | 189 | 370 | 827
6 | 268 | 189 | 404 | 861
7 | 268 | 294 | 403 | 965
8 | 268 | 294 | 365 | 927
Table 1. Path delays from signal input to flip-flops

## Slide 21: Post-synthesis and implementation
Resource | Utilization | Available | Utilization %
LUT | 51497 | 331680 | 15.53
LUTRAM | 312 | 146880 | 0.21
FF | 122876 | 663360 | 18.52
BRAM | 166 | 1080 | 15.37
IO | 319 | 520 | 61.35
BUFG | 14 | 624 | 2.24
MMCM | 2 | 12 | 16.67
Table 2. The summary of resource utilization.

## Slide 22: Result
Temperature and time resolution over 14-day continuous operation.
The standard deviations of the time bin width for each TDC channel.

## Slide 23: Performance with RPC detector: results
ch 0-7
ch 48-55
ch 72-79
Hit map matches the setup, with scintillators placed at the indicated positions.
Leading time resolution are less than 100ps.
Pulse widths are 5–10 ns and the cluster-size mean is 1.24.