## Slide 1: ORBIT FEEDBACK SYSTEMS and XBPMs
J-C. Denard, N. Hubert * and L. Cassinari*
* Diagnostics Group
Synchrotron SOLEIL, France
On behalf of the OFB team
Workshop on Accelerator R&D for Ultimate Storage Rings
30 October – 01 November 2012
Beijing

## Slide 2: Outline
Orbit Stability
  USR stability requirements
  Perturbation sources
  SOLEIL design for stability
BPM system for USR
  BPM supports
  Vacuum issues
  Electronics
Orbit Feedback systems
  Feedback Bandwidth
  Short term stability
  Slow drifts
XBPMs and photon beam feedback systems
  Existing XBPMs in e-beam FB loops
  XBPM developments needed
  Other kind of photon beam feedback systems

## Slide 3: USR Stability Requirements
The minimum attainable beam size is the diffraction limited photon beam
Photon beam dimensions from an undulator:
Chart: 
| 1 | 2 | 5 | 10 | 20 | 50 | 100
 | 5.6044 | 3.9629 | 2.5064 | 1.7723 | 1.2532 | 0.79259 | 0.56044
 | 17.6068 | 12.4499 | 7.874 | 5.5678 | 3.937 | 2.49 | 1.7607
E (keV)
10 keV photons
w/ 2m undulator
Size σ = 1.8 µm
Div. σ’ = 5.6 µrad
σ/10 requirements
∆ position  180 nm
∆ angle  560 nrad

## Slide 4: Sources of Perturbations (4)
Thermal effects
Insertion Devices
Cycling Booster
operation
Experimental hall activities
Ground vibrations
mains +harmonics
Frequency (Hz)
1
10
102
103
10-1
10-2
Time Period (s)
10-2
10-1
1
102
10
10-3
N
O
I
S
E
S
O
U
R
C
E
S
An orbit feedback is needed to stabilize the beam position from DC up to ~200 Hz
  Beamline Acquisition Time:
    If TPERTURBATION > TINTEGRATION → Emittance growth
    If TPERTURBATION < TINTEGRATION → Drifts
1/Ti
Drift
Emittance growth only

## Slide 5: Soleil Design for Stability (1)
Machine & expérimental hall on a common concrete slab made on top of 11 to 17 m long pilars resting on compact « Fontainebleau » sand.
Building structure and crane supports are decoupled from the stable slab.
Machine tunnel air and machine water cooling stabilized to 21±0.1°C. Actual air temperature is not exactly at that temperature. Stabilization is closer to ±0.25 °C around average values that can be different from 21°C. The important thing is to keep the temperatures stable whatever they are.
Girder rigidity in order to have all resonances above ground excitation frequencies (> 50 Hz)
The BPMs must be stable (beam stability cannot be better than that of the BPM system.)
  The BPM is the fixed point of its vacuum chamber section; the rest of vacuum chamber can move during beam delivery and during baking. There is a bellow on at least one side of each BPM. Bellows are shielded for low vacuum chamber impedance
  The BPM supports are rigid in transverse directions, the vertical stability of 4 BPMs (for the long beamlines) is garanteed by INVAR supports.
  The BPM blocks are water cooled.

## Slide 6: Soleil Design for Stability (2)
The golden orbit is done with respect to the quadrupole magnetic center (via Beam Based Alignment). But after that, the reference becomes the concrete slab via the BPMs adjacent to each straight section. With a perfect orbit feedback algorithm, only slab deformation could spoil beam stability.
Top up operation
Feedforward corrections for changes of undulator gap and phase, to better than 10 µm (should be down to 2 µm for USRs)
Booster compensation cables have been installed for correcting a 3 Hz spectrum line due to Booster ramping power supply cables. Cooling fans have been displaced for suppressing 47 to 108 Hz spectrum lines.
Orbit feedback systems. This is the last stone of the stability building blocks. It suppresses the ID feedforward residuals, small temperature variation effects in the tunnel, but not large ones. The FOFB is not very efficient above 70 Hz (with a 200 Hz BW).

## Slide 7: Outline
Orbit Stability
  USR stability requirements
  Perturbation sources
  SOLEIL design for stability
BPM system for USR
  BPM supports
  Vacuum issues
  Electronics
Orbit Feedback systems
  Feedback Bandwidth
  Short term stability
  Slow drifts
XBPMs and photon beam feedback systems
  Existing XBPMs in e-beam FB loops
  XBPM developments needed
  Other kind of photon beam feedback systems

## Slide 8: BPM Mechanical and Vacuum Design
Mechanical design:
  Thermal expansion coefficients of steel is 11ppm/°C => at 1.2 m height and 0.1°C drift, one gets 1.3 µm drift.
  INVAR is ~10 times better. A solution could be Invar BPM supports, at least at the critical locations.
Vacuum BPM Design:
  With 20 mm diameter vacuum chambers, the button BPMs might loose too much sensitivity
  New BPM design for small round pipe
  Shielded bellows matching the small round tubes.

## Slide 9: BPM Electronics for USRs
The Libera commercialized by Instrumentation Technologies is the state of the art.
Libera fulfills the USR requirements:
  Low temperature drifts
  Low beam current dependence (< 200 nm in top-up mode)
  Low bunch pattern dependence
  Position interlock built in
  Fast orbit feedback features built in (for BW up to ~200 Hz; it is possible to go up to 500 Hz, like at Petra III)
  Post mortem data available
  High reliability
(thanks to the electrode switching
scheme)

## Slide 10: USR Orbit Feedback Systems
Probably similar to existing systems
  Bandwidth extension to 500 Hz would supress better the 50Hz spectrum lines of the mains (60 Hz in Americas)
  Vacuum chamber space with thin stainless steel walls need to be reserved for fast correctors
  If a square correction matrix is not possible, the correction algorithm should favor the few beamlines with tightest stability requirements

## Slide 11: SOLEIL Feedback Characteristics
Parameters (H,V) | SOFB | FOFB
BPM # | 122, 122 | 122, 122
Corrector # | 57, 57 arcs | 50, 50 Straight sections
Sing. Value # | 57, 57 | 46, 47
Corrector max. | 1.0, 0.7 mrad | 28, 23 µrad
Correction rate | 0.1 Hz | 10 kHz
Bandwidth | 0-0.05 Hz | 0-200 Hz
Efficiency | IDs and arcs | Mostly IDs

## Slide 12: Slow orbit feedback system
Orbit stability requirements
  1/10th of beam sizes (H: 20 µm, V: 0.8 µm)
122 BPMs
  50 with supports on tunnel slab (4 of them on INVAR stands)
  72 with supports bolted on top of the girders
  Libera Electronics noise: 0.02 µm rms (0.1-10 Hz)
57 Horizontal and vertical steerers
  Cut-off frequency a few 10’s of Hz (Al Chamber)
RF-frequency Algorithm: SVD 57 singular values in both plane
  correction rate: every 10 s
  5 s averages of 10Hz BPM data
  RF frequency correction step size : 0.3 Hz
Matlab based control application
Libera module
Steerers are made with secondary coils in sextupole magnets

## Slide 13: Fast Orbit Feedback System
122 LIBERA modules
100 power-supplies
122 BPMs
50 fast correctors
An all-embedded and distributed solution
  Diamond Communication Controller for data distribution
  Computation done inside Libera FPGA (1 H&V corrector ↔ 1 Libera)
  Corrector power supplies controlled by Libera modules
200 Hz
0 dB gain

## Slide 14: Both Slow and Fast Systems work together
The Soleil orbit feedback systems are unique so far : they have two distinct sets of correctors for the slow and the fast system.
Advantages: air correctors over a thin stainless steel pipe can make fast corrections.
An algorithm periodically unloads the fast (and weak) correctors onto the slow (and more powerful) correctors.
The slow feedback benefits from its larger number of correctors for better correcting slow drifts.
It looks like the best way to go for future USRs.

## Slide 15
Vertical beam motion averaged on all e-BPMs ~300 nm (0.1-500 Hz )
It means 200 nm RMS in the middle of the straight sections
Vertical plane
FOFB ON
FOFB OFF
FOFB ON
FOFB OFF
SOLEIL Fast Orbit Feedback System

## Slide 16: Fast Orbit Feedback Bandwidth
SR Facility | FB type (users operation) | Sets of correctors | Bandwidth
ALBA* | Fast | Strong | DC - 130 Hz
BESSY II * | Fast | Strong | DC - 40 Hz
DIAMOND | Fast | Strong | DC - 130 Hz
ELETTRA | Fast | Strong | DC - 150 Hz
ESRF-U* | Fast | Strong | DC - 150 Hz
SLS | Fast | Strong | DC - 100 Hz
SPEAR3 | Fast | Strong | DC - 100 Hz
ALS | Slow + Fast | Strong ( + Fast as a subset of slow ones) | DC - 60 Hz
APS | Slow + Fast | Strong (+ Fast as a subset of slow ones) | DC - 100 Hz
ESRF | Slow + Fast | Strong + Fast | DC - 150 Hz
NSLS II* | Slow + Fast | Strong + Fast | DC - 500 Hz
PETRA III | Slow + Fast or Fast | Strong + Fast | Dead-band or DC - 500 Hz
SOLEIL | Slow + Fast | Strong + Fast | DC - 20 Hz
SSRF | Slow + Fast | Strong + Fast | DC - 100 Hz
* Feedback system not yet commissioned
N. Hubert
Bandwidth Limitations:
BPM electronics latency
Correctors & Power supplies
Vacuum Chamber Eddy Currents
Fast network Latency
SOLEIL noise suppression requires BW > 150 Hz
No real need for BW > 500 Hz

## Slide 17: Slow Drifts
1 week Beam Delivery
1 µm or µrad
Horizontal plane
1 µm or µrad
Vertical plane
X
X’
Y
Y’
NanoTomography source point position and angle during 1 week
1-week rms Drifts
∆X = 0.517 µm rms
∆X’ = 0.140 µm rms
∆Y = 0.134 µm rms
∆Y’ = 0.130 µm rms
peak to peak values
∆X = 5.16 µm p to p
∆X’ = 2.01 µm p to p
∆Y = 1.36 µm p to p
∆Y’ = 1.05 µm p to p

## Slide 18: Outline
Orbit Stability
  USR stability requirements
  Perturbation sources
  SOLEIL design for stability
BPM system for USR
  BPM supports
  Vacuum issues
  Electronics
Orbit Feedback systems
  Feedback Bandwidth
  Short term stability
  Slow drifts
XBPMs and photon beam feedback systems
  Existing XBPMs in e-beam FB loops
  XBPM developments needed
  Other kind of photon beam feedback systems

## Slide 19: Photon BPMs of BM Beamlines in OFB at SOLEIL
Introduction of photon XBPMs in the orbit feedback loop
  Stabilization of the Photon beam position on Bending Magnet Beamlines
XBPM in the global loop
  XBPM is considered as an e-BPM
  Possible only if XBPM measurement is reliable at any time (no gap/phase dependence)
Fully tested; will be in routine operation in November 2012
Beamline | RMS drift w/o XBPMs | RMS drift w/ XBPMs | Stability Gain
ODE | 1.25 µm rms | 0.37 µm rms | 3.3
METRO | 0.46 µm rms | 0.24 µm rms | 1.9
SAMBA | 0.38 µm rms | 0.23 µm rms | 1.7
DIFFABS | 0.33 µm rms | 0.16 µm rms | 2.1
Variation (RMS) of the vertical position on x-BPM N°2 with and without x-BPM N°1 in the slow feedback loop (experiment duration = 6 hours).
Vertical beam position measured on second x-BPMs (not included in the feedback loop) at the METRO, SAMBA and DIFFABS bending magnet beamlines. First x-BPM is alternatively included or excluded from the slow feedback loop. Improvement on vertical stability can be seen when the SOFB runs with the x-BPMs

## Slide 20: XBPMs in Slow OFB at Diamond
Without XBPM feedback.. ..with XBPM feedback.
Front End XBPMs
Beamline slit currents
(Data from beamline I19)
Courtesy C. Bloomer

## Slide 21: XBPMs, State of the Art
There are Reliable XBPMs; maybe not at the 200 nm level yet.
  BM beamline XBPMs (blade type) OK in vertical plane only
  ID XBPMs (blades type) OK in both planes at fixed gap.
  Most XBPMs are in stainless steel and are sensitive to temperature variations.
Present Development SOLEIL/LNLS of a stable blade XBPM at the 200 nm level, even after a beam loss. Not yet tested.

## Slide 22: XBPM R&D for USRs
Need hard X-ray photon BPMs stable with gap changes. Not an easy tasks
Need soft X-ray XBPMS stable with gap/phase changes. This is even more difficult
A better way might be to maximize the white beam flux through the beamline like it is done in some beamlines after the monochromator (exemple at DLS, Bloomer SRI2012)
One should be able to do better by improving communication between machine and beamlines.
Then one could install a sensor in the Beamline for a feedback on the electron beam.

## Slide 23
Using this system the beam intensity at the sample point can be maintained for several hours to 0.1% r.m.s., an improvement of a factor of 5.
Using DCM pitch modulation and feedback to improve long term X-ray beam stability
SRI2012, courtesy of C Bloomer (DLS)

## Slide 24: White Beam XBPM: Tentative Requirements for USRs
ID Gap and Phase independence
Beam current independence (for all XBPM current range)
Stability  10 µm during acquisition time for XBPM performing 560 nrad angular control at 20 m from the source point.
But in order to measure the real machine stability with an instrument outside the orbit feedback loop one needs better than 200 nm resolution and stability like for eBPMs.
Stability should be evaluated in the worse case : Difference (before – after) a beam loss at full current. It takes thermal inertia into account.

## Slide 25: XBPM for checking Machine Stability
upstream
e-BPM
downstream
e-BPM
X-BPM
# 2
slits
~3 m
~10 m
~10 m
Concrete slab
HLS
undulator
Ring
BL Frontend
Beamline
X-BPM
#1
The second XBPM for Feedback; the first one outside the FB loop monitors machine
stability
The HLS monitors slab deformation; it could eventually correct for it via a dedicated
slow feedback

## Slide 26: XBPM electronics Requirements
Stable within its environment temperature range
Compatible with orbit feedback systems (slow and fast). A « photon-Libera » is commercialized (IT/Soleil collaboration) with these specifications
Beam current dependence (measured with a simulated 0.5 mm off-centered beam) over the useful beam current dynamic range (includes all XBPMs, BM and IDs)