## Slide 1: The Comparison of API T3 and Leica AT401 in the Test and Survey
CAI .Guozhu
MAN. Kaidi
Institute of Modern Physcis,CAS
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 2: Outline
Different techniques of distance and angle;
Metrology results;
Survey results;
Conclusion
The 13th international Workshops on Accelerator Alignment (IWAA),HEP,2014

## Slide 3: 1.Different techniques of distance and angle
Leica AT401 API T3
Distance measurement with laser Interferometer Mode and ADM
Absolute Angle Encoders
ADM
Relative Angle Encoders
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014
The tracker is the spherical measurement systemt with two angles and one distance,but…

## Slide 4
Typical Laser Tracker:
For example, Angle accuracy VS Distance accuracy at 10m length
Distance accuracy(IFM ): 0.5m/m, 5m in 10m;
Angle accuracy : 1 arcsecond,48 m in 10m.
Above results in 2sigma
ρ=206265
>>
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 5: 2.Calibration in CHANGCHENG Institute of Metrology & Measurement
Dual-frequency interferometer uncertainty：1X10-8，k=2
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014
The homodromous and vertical moving the target with the interferometer was designed to calibrate distance and angle accuracy on the testing bench.

## Slide 6: Leica AT401 API T3
Chart: 
| 1.032465 | 2.1086140000000002 | 3.1081920000000012 | 4.0464680000000124 | 5.1175959999999252 | 6.1798980000000014 | 7.1698159999999334 | 8.015369999999999 | 9.0816990000000004 | 10.039658000000001 | 11.026895 | 12.021499 | 12.983865 | 14.033735 | 14.959294000000074 | 15.964804000000004 | 17.046525999999989 | 18.079253000000001 | 19.020554000000001 | 20.033498999999999 | 21.043274 | 22.033929000000001 | 23.092929999999889 | 24.163464999999999 | 25.085691999999753 | 26 | 27 | 28 | 29 | 30
AT401 | -0.034 | -0.037 | 0.055 | 0.076 | 0.051 | 0.083 | 0.039 | 0.06 | 0.05 | 0.05 | 0.039 | 0.052 | 0.077 | 0.068 | 0.064 | 0.072 | 0.069 | 0.057 | 0.049 | 0.063 | 0.057 | 0.047 | -0.076 | -0.034 | -0.068 | -0.079 | -0.084
T3 | -0.002 | 0.001 | 0.001 | 0.002 | 0.003 | 0.002 | 0.002 | 0.003 | 0.004 | 0.003 | 0.001 | 0.001 | 0 | -0.001 | 0.002 | 0.003 | 0.003 | 0.004 | 0.004 | 0.005 | 0.005 | 0.006 | 0.006 | 0.006 | 0.007 | 0.007 | 0.008 | 0.007 | 0.009 | 0.01
Chart: 
| 1.0257559999999999 | 2.0053900000000002 | 3.0349529999999967 | 4.0300900000000004 | 5.0833930000000134 | 6.0207109999999755 | 7.0499020000000003 | 8.0248060000000017 | 9.0217779999999994 | 10.032134000000006 | 11.043236 | 12.127118999999999 | 13.046902999999999 | 14.055206000000076 | 15.066829 | 16.086181 | 17.086025000000003 | 18.088112999999673 | 19.128899000000001 | 20.208368 | 21.121697000000001 | 22.084099999999989 | 23.203259999999986 | 24.22240999999968 | 25.241575000000001 | 26.260739999999583 | 27.279900000000001 | 28.299060000000001 | 29.318219000000003 | 30.337379000000031
AT401 | 0 | -0.023 | -0.009 | 0.02 | 0.008 | 0.016 | 0.006 | -0.02 | 0.001 | 0 | -0.013 | -0.02 | -0.004 | 0.013 | 0.034 | 0.058 | 0.046 | 0.021 | 0.031 | 0.003 | 0.008 | 0.03 | 0.044 | 0.055 | 0.082 | 0.072 | 0.085 | 0.104 | 0.096 | 0.109
T3 | -0.01 | -0.007 | -0.013 | -0.003 | -0.01 | -0.017 | -0.013 | -0.02 | -0.03 | -0.043 | -0.059 | -0.045 | -0.041 | -0.035 | -0.029 | -0.032 | -0.029 | -0.019 | -0.014 | -0.011 | -0.016 | -0.007 | -0.004 | 0.009 | 0.007 | 0.006 | 0.009 | 0.012 | 0.017 | 0.02
m
mm
Comparison in Distance
Comparison in Angle
API T3
Leica AT401
VS
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 7: Layout of the Hirfl-CSR(Heavy ion research facility in lanzhou-Cooling Storage Ring)
(128.8m)
3.Survey in the CSR
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 8
Survey in the CSRm
The CSRm control network with the total 345 points in 15 stations one by one, used bundle adjustment to link the both adjacent stations and closed loop.
Instrument:API T3，temperature：26°C，
Time：2010.8.6
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 9: Analysis of instrument residual errors in SA
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014
Horizontal Angle is 0.6 arcsecond
Vertical Angle is 0.57 arcsecond
Distance is 0.017mm

## Slide 10
Ins | Ux | Uy | Uz | URx | URy | URz
T0 | 0.023 | 0.043 | 0.120 | 2.88 | 1.44 | 0.72
T1 | 0.025 | 0.050 | 0.117 | 2.88 | 1.44 | 0.72
T2 | 0.041 | 0.049 | 0.094 | 2.52 | 2.16 | 1.08
T3 | 0.048 | 0.035 | 0.092 | 1.8 | 2.52 | 0.72
T4 | 0.051 | 0.022 | 0.136 | 1.44 | 2.88 | 0.72
T5 | 0.043 | 0.027 | 0.118 | 1.8 | 2.88 | 0.72
T6 | 0.038 | 0.039 | 0.087 | 2.52 | 2.88 | 0.72
T7 | 0.027 | 0.048 | 0.110 | 2.52 | 1.8 | 0.72
T8 | 0.028 | 0.048 | 0.111 | 2.88 | 1.8 | 0.72
T9 | 0.025 | 0.049 | 0.135 | 2.52 | 1.8 | 0.72
T10 | 0.040 | 0.048 | 0.097 | 2.52 | 2.52 | 0.72
T11 | 0.051 | 0.040 | 0.094 | 2.16 | 2.88 | 0.72
T12 | 0.051 | 0.026 | 0.167 | 1.44 | 2.88 | 0.72
T13 | 0.057 | 0.028 | 0.167 | 1.8 | 2.88 | 1.08
T14 | 0.035 | 0.043 | 0.108 | 2.88 | 1.8 | 0.72
Location uncertainty(6 parameters) in the network(T3).
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014
The statistic results of all the instruments shows the Uz,URx,Ury is bigger than other, the reason maybe is the shape of the network or the performance of the tracker.

## Slide 11: Survey in the CSRe with AT401
The CSRe control network with 10 stations one by one, and closed loop too.
Time：2013.8
Instument:AT401，temperature：25°C，
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 12: Location uncertainty in the network(AT401 in survey CSRe).
Inst | Ux | Uy | Uz | URx | URy | URz
T0 | 0.01 | 0.02 | 0.05 | 1.08 | 1.08 | 0.36
T1 | 0.01 | 0.02 | 0.06 | 1.08 | 1.44 | 0.36
T2 | 0.03 | 0.02 | 0.04 | 1.08 | 1.8 | 0.36
T3 | 0.03 | 0.02 | 0.06 | 1.08 | 2.52 | 0.72
T4 | 0.02 | 0.02 | 0.04 | 2.16 | 1.44 | 0.72
T5 | 0.01 | 0.02 | 0.07 | 2.52 | 1.44 | 0.72
T6 | 0.01 | 0.02 | 0.07 | 1.08 | 1.44 | 0.36
T7 | 0.02 | 0.01 | 0.03 | 0.72 | 1.44 | 0.36
T8 | 0.02 | 0.02 | 0.05 | 0.72 | 1.44 | 0.36
The results are same as the survey in the CSRm with T3.
So it is the weakness in height measurement by tracker in the ring survey network,leveling is necessary was proved again!
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 13: All previous CSRm survey list:
Year | Inst | Stations | Points | Overall error(mm) | Up-mean(mm) | Uh(“) | Uv（“） | Ud（mm）
2010.8 | T3 | 15 | 345 | 0.03 | 0.13 | 0.6 | 0.57 | 0.017
2011.4 | T3 | 15 | 320 | 0.03 | 0.17 | 0.63 | 0.63 | 0.020
2011.8 | AT401 | 13 | 270 | 0.047 | 0.25 | 1.5 | 0.49 | 0.038
2012.8 | T3 | 14 | 297 | 0.067 | 0.20 | 1.12 | 1.16 | 0.076
2013.7 | AT401 | 14 | 356 | 0.07 | 0.12 | 2.0 | 0.6 | 0.038
2013.8 | T3 | 14 | 393 | 0.06 | 0.07 | 0.8 | 1.28 | 0.032
1,Overall error <=0.07mm,and uncertainty average is about 0.2mm;
2,The performance of AT401 is stable, althought the distance performance is worse than T3;
3,The stablity of T3 needs to be pay attention to do field check and calibration need to be done usually!
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 14
4.Conclusion
T3 is outstanding distance performance along the laser emission direction.
AT 401 is the equivalent of 3D performance with T3 in the network surveying.
Both AT 401 and T3 is the low accuracy in height measurement of the ring control network.
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014

## Slide 15
Thanks for attention very much!
Lanzhou City
The 13th international Workshops on Accelerator Alignment (IWAA),IHEP,2014