## Slide 1: The Comparison Of API T3 And Leica AT401 in the Test and Survey
Cai guozhu
MAN. Kaidi
Institute of Modern Physcis,CAS

## Slide 2: Outline
Difference technolgy of distance and angle;
Metrology results;
Survey results;
Conclusion

## Slide 3: 1.Difference technology of distance and angle
Leica AT401 API T3
Distance measurement with laser Interferometer Mode and ADM
Absolute Angle Encoders
ADM
Relative Angle Encoders

## Slide 4
Typical Laser Tracker:
For example, at 10m Angle accuracy VS Distance accuracy
Distance accuracy(IFM ): 0.5m/m, 5m in 10m;
Angle accuracy : 1 arcsecond,48 m in 10m.
Above results in 2sigma
ρ=206265
>>

## Slide 5: 2.Calibration in CHANGCHENG Institute of Metrology & Measurement
Dual-frequency interferometer uncertainty：1X10-8，k=2

## Slide 6: Leica AT401 API T3
Chart: 
| 1.032465 | 2.1086140000000002 | 3.1081920000000012 | 4.0464680000000124 | 5.1175959999999296 | 6.1798980000000014 | 7.1698159999999378 | 8.0153700000000008 | 9.0816990000000004 | 10.039658000000001 | 11.026895 | 12.021499 | 12.983865 | 14.033735 | 14.959294000000074 | 15.964804000000004 | 17.046525999999989 | 18.079253000000001 | 19.020554000000001 | 20.033498999999999 | 21.043274 | 22.033929000000001 | 23.092929999999889 | 24.163464999999999 | 25.085691999999771 | 26 | 27 | 28 | 29 | 30
AT401 | -0.034 | -0.037 | 0.055 | 0.076 | 0.051 | 0.083 | 0.039 | 0.06 | 0.05 | 0.05 | 0.039 | 0.052 | 0.077 | 0.068 | 0.064 | 0.072 | 0.069 | 0.057 | 0.049 | 0.063 | 0.057 | 0.047 | -0.076 | -0.034 | -0.068 | -0.079 | -0.084
T3 | -0.002 | 0.001 | 0.001 | 0.002 | 0.003 | 0.002 | 0.002 | 0.003 | 0.004 | 0.003 | 0.001 | 0.001 | 0 | -0.001 | 0.002 | 0.003 | 0.003 | 0.004 | 0.004 | 0.005 | 0.005 | 0.006 | 0.006 | 0.006 | 0.007 | 0.007 | 0.008 | 0.007 | 0.009 | 0.01
Chart: 
| 1.0257559999999999 | 2.0053900000000002 | 3.0349529999999967 | 4.0300900000000004 | 5.0833930000000134 | 6.0207109999999755 | 7.0499020000000003 | 8.0248060000000017 | 9.0217779999999994 | 10.032134000000006 | 11.043236 | 12.127118999999999 | 13.046903 | 14.055206000000076 | 15.066829 | 16.086181 | 17.086025000000003 | 18.088112999999698 | 19.128899000000001 | 20.208368 | 21.121697000000001 | 22.084099999999989 | 23.203259999999986 | 24.222409999999709 | 25.241575000000001 | 26.260739999999618 | 27.279900000000001 | 28.299060000000001 | 29.318219000000003 | 30.337379000000031
AT401 | 0 | -0.023 | -0.009 | 0.02 | 0.008 | 0.016 | 0.006 | -0.02 | 0.001 | 0 | -0.013 | -0.02 | -0.004 | 0.013 | 0.034 | 0.058 | 0.046 | 0.021 | 0.031 | 0.003 | 0.008 | 0.03 | 0.044 | 0.055 | 0.082 | 0.072 | 0.085 | 0.104 | 0.096 | 0.109
T3 | -0.01 | -0.007 | -0.013 | -0.003 | -0.01 | -0.017 | -0.013 | -0.02 | -0.03 | -0.043 | -0.059 | -0.045 | -0.041 | -0.035 | -0.029 | -0.032 | -0.029 | -0.019 | -0.014 | -0.011 | -0.016 | -0.007 | -0.004 | 0.009 | 0.007 | 0.006 | 0.009 | 0.012 | 0.017 | 0.02
m
mm
Comparison in Distance
Comparison in Angle
API T3
Leica AT401
VS

## Slide 7: Layout of the CSR(Cooling Storage Ring) in IMP
(128.8m)
3.Survey in the CSR

## Slide 8
Survey in the CSRm
Time：2010.8.6
Instument:API T3，temperature：26°C，
15 stations.

## Slide 9: Analysis of instrument residual errors in SA

## Slide 10
Ins | Ux | Uy | Uz | URx | URy | URz
T0 | 0.023 | 0.043 | 0.120 | 2.88 | 1.44 | 0.72
T1 | 0.025 | 0.050 | 0.117 | 2.88 | 1.44 | 0.72
T2 | 0.041 | 0.049 | 0.094 | 2.52 | 2.16 | 1.08
T3 | 0.048 | 0.035 | 0.092 | 1.8 | 2.52 | 0.72
T4 | 0.051 | 0.022 | 0.136 | 1.44 | 2.88 | 0.72
T5 | 0.043 | 0.027 | 0.118 | 1.8 | 2.88 | 0.72
T6 | 0.038 | 0.039 | 0.087 | 2.52 | 2.88 | 0.72
T7 | 0.027 | 0.048 | 0.110 | 2.52 | 1.8 | 0.72
T8 | 0.028 | 0.048 | 0.111 | 2.88 | 1.8 | 0.72
T9 | 0.025 | 0.049 | 0.135 | 2.52 | 1.8 | 0.72
T10 | 0.040 | 0.048 | 0.097 | 2.52 | 2.52 | 0.72
T11 | 0.051 | 0.040 | 0.094 | 2.16 | 2.88 | 0.72
T12 | 0.051 | 0.026 | 0.167 | 1.44 | 2.88 | 0.72
T13 | 0.057 | 0.028 | 0.167 | 1.8 | 2.88 | 1.08
T14 | 0.035 | 0.043 | 0.108 | 2.88 | 1.8 | 0.72
Location uncertainty in the network(T3).
Locating instruments with 6 parameter.

## Slide 11: Survey in the CSRe with AT401
Time：2013.8
Instument:AT401，temperature：25°C，
10 stations.

## Slide 12: Location uncertainty in the network(AT401 in survey CSRe).
Inst | Ux | Uy | Uz | URx | URy | URz
T0 | 0.01 | 0.02 | 0.05 | 1.08 | 1.08 | 0.36
T1 | 0.01 | 0.02 | 0.06 | 1.08 | 1.44 | 0.36
T2 | 0.03 | 0.02 | 0.04 | 1.08 | 1.8 | 0.36
T3 | 0.03 | 0.02 | 0.06 | 1.08 | 2.52 | 0.72
T4 | 0.02 | 0.02 | 0.04 | 2.16 | 1.44 | 0.72
T5 | 0.01 | 0.02 | 0.07 | 2.52 | 1.44 | 0.72
T6 | 0.01 | 0.02 | 0.07 | 1.08 | 1.44 | 0.36
T7 | 0.02 | 0.01 | 0.03 | 0.72 | 1.44 | 0.36
T8 | 0.02 | 0.02 | 0.05 | 0.72 | 1.44 | 0.36
The results are same as the survey in the CSRm with T3.

## Slide 13: All previous CSRm survey list:
Year | Inst | Stations | Points | Overall error(mm) | Up-mean(mm) | Uh(“) | Uv（“） | Ud（mm）
2010.8 | T3 | 15 | 345 | 0.03 | 0.13 | 0.6 | 0.57 | 0.017
2011.4 | T3 | 15 | 320 | 0.03 | 0.17 | 0.63 | 0.63 | 0.020
2011.8 | AT401 | 13 | 270 | 0.047 | 0.25 | 1.5 | 0.49 | 0.038
2012.8 | T3 | 14 | 297 | 0.067 | 0.20 | 1.12 | 1.16 | 0.076
2013.7 | AT401 | 14 | 356 | 0.07 | 0.12 | 2.0 | 0.6 | 0.038
2013.8 | T3 | 14 | 393 | 0.06 | 0.07 | 0.8 | 1.28 | 0.032
1,Overall error <=0.07mm,and uncertainty average is about 0.2mm;
2,The performance of AT401 is stable, althought the distance performance is worse than T3;
3,The stablity of T3 needs to be pay attention, the field check and calibration need to be done usually!

## Slide 14: 4.Conclusion
T3 is outstanding distance performance along the laser emission direction.
AT 401 is the equivalent of 3D performance with T3 in the network surveying.
AT 401 is more reliable performance than T3,and more portable in the survey.

## Slide 15
Thanks for attention very much!