## Slide 1: IHEP 9-cell Cavity Work (Oct. 31 - Nov. 4)
Jiyuan Zhai, Conglai Yang,
Motoaki Sawabe, Mineyuki Asano
November 7, 2016, KEK, Japan

## Slide 2: Weekly Schedule
Date | Work
Mon (10-31) | Pre-EP setup
Tue (11-1) | EP facility trouble (reservoir acid temperature monitor communication problem; fixed by changing the monitor-box with the one in the 2nd EP system)
Wed (11-2) | Pre-EP 5 μm, EP1 setup (one day delay)
Thu (11-3) | 文化の日
Fri (11-4) | EP1 100 μm, ultrasonic cleaning, HPR, dry

## Slide 3: EP Facility

## Slide 4: Pre-EP Procedure
Time | Process | Details
09:05 | Pre-EP 5 um | 5 V, cavity rotate 1 rpm, 40 min 55 s; if > 40 C, power off, water shower cooling
10:01 | Surface stabilization | Rotate 3 rpm, 20 min
10:21 | EP acid drain and water rinsing | 5 min vertical, 2 min horizontal, 2 min vertical; 10 min water in (continuous), 3 min drain, about 1 h (5 times)
11:45 | Cavity and EP components washing | cathode rod, flanges, o-rings etc.
18:30 | EP1 setup | temperature sensor, hook

## Slide 5: Pre-EP Parameters

## Slide 6: EP Support Jig

## Slide 7: Pre-EP Setup

## Slide 8: Pre-EP

## Slide 9: Cathode and EP Flange Cleaning

## Slide 10: EP1 Procedure
Time | Process | Details
08:57 | EP1 (bulk EP) 100 um | 25 V, ~ 50 mA/cm2, 4 h 21 min 55 s
13:19 | Surface stabilization | Rotate 3 rpm, 20 min
13:39 | EP acid drain and water rinsing | 5 min vertical, 2 min horizontal, 2 min vertical; 10 min water in (continuous), 3 min drain, about 1 h (5 times)
15:12 | Cavity and EP components washing | cathode rod, flanges, o-rings etc.
16:29 | Ultrasonic cleaning | FM-20 2% ultrasonic 50 C, 15 min
16:44 | Water washing | 
17:26 | HPR (1 h) | cavity inside, 8 MPa, 10 rpm, 200 mm/min
18:30 | Dry | 

## Slide 11: EP1 Parameters

## Slide 12: EP1 Parameters

## Slide 13
Downstream beam pipe temperature high. Cathode mask problem?
Not critical for EP1, but EP2 should be careful.

## Slide 14: Cathode Insert

## Slide 15: EP1 Setup

## Slide 16: Cavity Temperature Monitor

## Slide 17: EP1

## Slide 19: HPR and Dry