## Slide 1: HTS Magnet R&D at BNL
Ramesh Gupta
November 17, 2017

## Slide 2: Recent HTS Magnet Programs at BNL
Demonstration of HTS/LTS hybrid common coil dipole
  Use more expensive HTS where field is high and less expensive LTS (Nb3Sn and/or NbTi) where field is low (part of PBL/BNL STTR program)
High Field (25 T), Large Aperture (100 mm) HTS Solenoid
  HTS subjected to high field and high stresses
  Part of “Dark Matter Axion Search Program” at IBS, Korea. This magnet must be a reliable user magnet
BNL has the largest and the most versatile HTS magnet program in any US lab

## Slide 3: Unique BNL Common Coil Dipole
Common coil 2-in-1 dipole design with simple racetrack coils that are common to both apertures
Record demonstration of “React & Wind” technology
Structure specifically designed to provide a large open space (31 mm wide, 338 mm high)
New racetrack coils can be inserted in the magnet without requiring any disassembly or reassembly
New insert coils become an integral part of the magnet. Coil tests become magnet tests
Rapid-turn-around (<2 years), lower cost approach (allowed HTS/LTS hybrid dipole (<2 years, <1M$)
Empty space

## Slide 4
Insert coils in Empty space
Design
A pair of HTS insert coils
HTS/LTS Hybrid Dipole Structure

## Slide 5: HTS Coil Winding (two coils wound)
Two coils used ~300 meters of 4 mm equivalent
Conductor:
12 mm ASC tape
Insulation:
Nomex

## Slide 6: Retest of Nb3Sn Common Coil Dipole After a Decade
Current (A)
~9.5T@10kA
Short Sample: 10.8 kA (reached during 2006 test)
Retest: No quench to 10 kA (>92% of short sample)
Worked well when tested again after a decade
Was a display piece of Nb3Sn “React & Wind” technology for dipole

## Slide 7: HTS/LTS Hybrid Dipole Quench Test Results

## Slide 8: Challenges with the HTS/LTS Hybrid Dipole
Large coupling between HTS/LTS coils
    Maximum current in HTS: ~800 A
    Maximum current in Nb3Sn: ~10 kA
Protection of the HTS coils at 4 K
Quench protection of HTS coils in HTS/LTS hybrid configuration
Questions:
  Can HTS coil survive quenches without significant degradation?
  Can HTS coils be operated like the LTS coils?

## Slide 9: HTS/LTS Hybrid Operation and Quench Protection
HTS & LTS powered separately
Common quench platform; fast energy extraction from both coils
Quench detection response time: < 5 msec
Coil current interruption: < 10 micro-second after detection
HTS coil shut-off: a few msec
High power IGBT switches
Electronic threshold for quench detection: ~100 micro-volts
HTS Quench threshold : 5 mV
HTS coils were actually tested in more brutal conditions: ~200 mV (like LTS coils)

## Slide 10
Chart: 
Coil Voltage (V)
HTS coils operated like LTS coils
Significant voltage in HTS coils: >0.2 Volts

## Slide 11
Chart: Hybrid Dipole Field (T)
| 5.477777777777777 | 6.79
 | 0 | 756
 | 0 | 741
 | 0 | 712
 | 0 | 682
 | 0 | 656
 | 0 | 800
 | 0 | 500
 | 0 | 635
 | 0 | 500
 | 0 | 500
 | 0 | 500
 | 0 | 500
Operation of HTS/LTS Hybrid Dipole
(HTS coils ramped-up in different fields of Nb3Sn coils)
Encouraging Results:
HTS coils were ramped to quench, just like LTS coils
No degradation in HTS coils despite a number of quenches
Significant demonstration. 8.7 T may be the highest field HTS/LTS hybrid dipole magnet
Performance limited by the leads (not by coils)
~14 T possible with new HTS tapes, in favorable direction

## Slide 12
High Field, Large Aperture HTS Solenoid for Axion Dark Matter Search

## Slide 13: Requirements for the IBS Solenoid
High Field : 25 T (must use HTS)
Large Volume: 100 mm bore, +/-100 mm long
      Stresses: J X B X R
Field quality: ~10%
Ramp-up time: up to 1 day
      Relaxed field quality and slow charging
User magnet: robust design, large Margin

## Slide 14
Field: 25 T@4 K
Bore: 100 mm
Stored Energy: 1.7 MJ
Hoop Stresses: 400 MPa
Conductor: HTS (2G)
HTS tape (ReBCO) used:
~6 km, 12 mm wide from SuperPower
Reached a critical field at 27 K: 12.5 T
  Expected field at 27 K
  New record for magnets operating over >10 K in similar solenoid
Test terminated after arcing in the leads
HTS Solenoid for SMES
BNL Experience with a Similar Solenoid

## Slide 15: Design Choices for IBS Solenoid
  No Insulation Coil (no between the turns)
    Bypasses current from one turn to the next turn in the event of a quench
    Takes benefit of the relaxed field quality and slow charging time
  Single Layer Design
    In “No Insulation” case, there could be an unbalanced force situation between the two layers when the current is bypassed
  Conductor: High Field, High Strength, 12 mm wide 2G ReBCO Tape
    50 mm Hastelloy, 20 mm copper, advanced pinning
  Critical Current Margin : ~50% @4K
    High performance at high fields
    Need ~10 km of 12 mm wide tape
  Stress/Strain Margin : ~50%
Healthy electrical and mechanical margin
Bmax: 25.2 T

## Slide 16: Mechanical Properties of the Conductor (1)
Courtesy: SuperPower
Requirement of Azimuthal stresses of ~500 MPa is met with 2G Tape having
50 micron Hastelloy
and 20 micron Copper
Transverse Load:
load on the wide face of conductor

## Slide 17
Requirement of ~200 MPa on the narrow face of the tape must be verified as no such is data available
Initial results are positive
Conductor must also meet high load (200 MPa) on the narrow side
New Apparatus to Apply 300 MPa Load Design
Mechanical Properties of the Conductor (2)

## Slide 18: No Insulation Coil Construction and Test Results

## Slide 19: No Insulation Double Pancake Coils
To obtain data and 4 K test experience with large “NI” coil early on, a coil wound with ~550 m of 12 mm wide ReBCO tape
i.d. = 100 mm
o.d. = 220 mm
Turns = 971
Significant instrumentation and 3 heaters for simulated defects

## Slide 20: Transfer Function Vs. Current in No Insulation Coils
Over Current
V-I curve
Current (A)
Reduction in Transfer Function implies that the effective number of turns are decreasing

## Slide 21: B, V and I in No Insulation Coil at Thermal Runaway
Thermal Runaway
B
I
V1
V2
Current limited by PS
Time (seconds)
Next: 4K measurements

## Slide 22: Summary
Encouraging Test Results of HTS/LTS Hybrid Dipole
Many LTS type quenches in HTS coils with no degradation
A robust rapid-turn-around, low-cost test facility
  Insert coils become an integral part of the magnet
High Field, large aperture HTS solenoid program gives an experience in dealing with the high stress HTS magnets