## Slide 1: STFC SRF test facility for ESS high beta cavities
CryoOps Workshop at IHEP
June 6th 2018
Louis Bizel-Bizellot
UKRI

## Slide 2: What is ESS ?
European Spallation Source
Currently under construction
User program for 2023
LINAC 100% functional in 2025
 LINAC which accelerate proton toward a tungsten target generating neutrons.
 2.86 ms long proton pulse at 2 GeV at 14 Hz
5 MW of average power with a DT = 4%

## Slide 3: Hβ Cavity Project at STFC
Production + qualification
x 84 Hβ SRF cavities
Assembled into 21 cryo-module by CEA

## Slide 4: Hβ Cavity Project at STFC
x 84
Ø 460 mm
1550 mm
Nb
Ti

## Slide 5: Hβ Cavity Project at STFC
Load 3 cavities
Helium connection
UHV connection
RF connection
1
2
3

## Slide 6: SRF cavity test – STFC facility
Cryostat
Outside Ø = 2.2 m
Height = 4.25 m
3 ESS high β cavities
2K box contains:
Cryogenic valves
2K heat exchanger
JT valve
Cryo-plant
Hélial 100
RF + Cryo Control center
4.2K LHe
50K GHe
Magnetic shield (µMetal) > 1 µT
2K pumps

## Slide 7: Magnetic shield

## Slide 8: Cryostat – the facility
Cavity 1
Cavity 2
Cavity 3
Mobile cleanroom for UHV connection
Valve
box
Cryostat

## Slide 9: Safety considerations – PRV
 | Loss of vacuum heat load | Ø safety valve required | Cross section burst disk required
Without MLI on cavities | 170 kW | 64 mm | 17680 mm²
With MLI on cavities | 42 kW | 32 mm | 4368 mm²
Low pressure helium guard (on 30 mbar circuit)
Pressure relief valves
1.35
bara
1.1
bara
Burst disk
1.45
bara
100 mbar GHe

## Slide 10: Safety Considerations
Multi-layer insulation ad-hoc suit design for repetitive installation. All in 1 piece

## Slide 11: Cryoplant
Helium liquefier
100 ltrs/h liquefaction rate
+
50 K He gas for radiation shield
Liquid helium Dewar
3000 ltrs

## Slide 12: Cryoplant – He & LN2 storage
400 Nm3
Recovery He
400 Nm3
Recovery He
640 Nm3
Pure He
3000 ltrs
Liquid He
15 m3
Liquid N2
Gas bag
40 bar
16 bar

## Slide 13: Cryoplant – Storage issues
Cold box
400 Nm3
Recovery He
400 Nm3
Recovery He
Gas bag
20 m3 at 40 bar
Dryer
640 Nm3
Pure He
(40 m3 at 16 bar)
Purifier
Doesn’t work bellow 20 bar
Add-on for an extra compressor
Use the recovery compressor

## Slide 14: Cryoplant – operation plan
Chart: Helium inventory within the test facility
| 0 | 12 | 16 | 24 | 34 | 39 | 48
Dewar LHe (l) | 2500 | 2000 | 1500 | 1000 | 500 | 500 | 2500
Cryostat LHe (l) | 0 | 0 | 500 | 250 | 500 | 0 | 0
Recovery He storage (bar) | 2 | 20.5 | 1.24 | 11.97 | 19.37 | 37.87 | 2
Pure He storage (bar) | 2 | 2 | 12 | 12 | 12 | 12 | 2
Cool down 4.2K
LHe fill in
Cool down 2K
Operation
Warm up & re-liquefaction
Dewar LHe (ltrs)
Recovery He storage (bar)
Pure He storage (bar)
Cryostat LHe (ltrs)

## Slide 15: Summary
AL Helial 100 fully commissioned and tested succesfully at Daresbury
Cryostat fully installed, commissioning on going
Improvement of the GHe management under study
Extra GHe storage solution under investigation
Thank you for your attention