## Slide 1: OPERATING STATUS OF BES SUPERCONDUCTING MAGNET
Zhongxiu Liu

## Slide 2: Outline
Brief introduction&Operating status
Faults
Aging Problems
Design of new valve box
Summary

## Slide 3: Brief introduction&Operating status
BES superconducting magnet(SSM) is currently the largest single superconducting magnet manufactured in China. The normal working current of the magnet is 3369 A, and the yoke iron provides the circuit of magnetic field, which provides a constant magnetic field with a center field strength of 1 T and a certain degree of uniformity for the Beijing spectrometer, so that the charged particles are deflected in the magnetic field. By measuring the tracks of charged particles in the main drift chamber, the momentum of the particles is obtained to research the properties of the basic particles.
Normal Run was from 2008 to now. SSM run about seven months every year.

## Slide 4
BES STRUCTURE

## Slide 5
During the past ten years,there were total 37 quenches.

## Slide 6: Current (2015 to 2016 run)

## Slide 7: Current (2016 to 2017 run)

## Slide 8: Current (2017 to 2018 run)

## Slide 9: Faults
Slow discharge during exciting.
Two slow discharges during exciting,and first discharge caused to quench. The reason for the two slow discharges is due to the aging of the device inside the quench detector. We tackled problem as the right diagrammatic sketch.
Qunech Detector
OMRON-PLC
OMRON-PLC
AD081
OC211
Channel 1
Channel 2
A-1
B-1
0
1
Current lead A
Current lead B
No.32 signal line—I7
No.32 signal line—I8
BES HALL
POWER CONTROLROOM
interlock PLC

## Slide 10: Aging Problems
Vacuum
Transition section in valve box
Design of new valve box

## Slide 11: Aging Problem－－Vacuum
From 2010 to now, the vacuum value kept rising from 2.2×10-2Pa to 2.7×10-2Pa, even to 3.8×10-2Pa in 2016,and back to 2.7×10-2Pa ,it was
not stable within last few years.
It increased significantly under 20K. We believe that it should be liquid helium leaked .
Chart: VACCUM-PI1563
| 39648 | 39697 | 39783 | 39825 | 39846 | 39863 | 39955 | 39996 | 40222 | 40259 | 40493 | 40525 | 40551 | 40581 | 40635 | 40669 | 40674 | 40704 | 40980 | 41000 | 41228 | 41315 | 41351 | 41583 | 41610 | 41613 | 41649 | 41701 | 41743 | 41786 | 41943 | 41991 | 41998 | 42025 | 42048 | 42088 | 42111 | 42320 | 42368 | 42389 | 42419 | 42478 | 42499 | 42751 | 42789 | 42818 | 42842 | 42872 | 43164 | 43246 | 43248
 | 3.79 | 3.56 | 3.65 | 3.93 | 3.42 | 3.67 | 3.44 | 3.65 | 2.31 | 2.36 | 2.18 | 2.22 | 2.12 | 2.16 | 2.2 | 2.18 | 2.17 | 2.11 | 2.11 | 2.11 | 2.2 | 2.07 | 2.07 | 2.93 | 2.88 | 2.93 | 2.95 | 2.88 | 2.93 | 2.8 | 2.63 | 2.39 | 2.66 | 2.62 | 2.67 | 2.64 | 2.69 | 3.45 | 3.65 | 3.76 | 3.75 | 3.77 | 3.88 | 2.71 | 2.65 | 2.7 | 2.69 | 2.61 | 2.8 | 3.2 | 2.9

## Slide 12
Aging Problem－－Transition section in valve box

## Slide 13
We worry about :
1.Temperature keep rising and out of control.
2.Burn the current transition section.
We did some experiments in Nov.12 2012 and adjusted warning threshold value of TSD-12 from 6.38 to 6.5
Aging Problem－－Transition section in valve box

## Slide 14
There are some bimetal joints in valve box. Now we have no methods to check these bimetal joints, the reason that the vacuum became worse is perhaps those joints leaking.
It had been leaked from the first cooling.
Aging Problem－－Bimetallic joints in valve box

## Slide 15: Aging Problem－－Transition section in valve box
From Feb.2010, the temperature of the transition section(TSD12) kept rising to 6.42K.

## Slide 16: Design of new valve box---Requirements and technical improvements
Technical improvements:
HTS current leads instead of gas-cooled current lead
Optimize the cryogenic pipeline, reduce number of the bimetal joints, to improve the vacuum
Optimization of cooling structure of the transition section
Requirements：
Unchange the overall dimensions of the valve box
Unchange the position and mode of the joint with the cryogenic system
Unchange position of the joints with power bus
Overall heat leakage can not be increased
Inside structure of the existing valve box
Design of the new valve box

## Slide 17: Design of new valve box- New current leads
Operating current: 3400A
Structure: binary current leads
Heat exchange section cooled by helium gas
HTS stack structure
Heat exchanger structure

## Slide 18: Design of new valve box- Cryogenic pipeline
Pipes from the coils are all aluminum tubes
There are several bimetal joints in the valve box, and these maybe the reason of the bad vacuum
Optimize the cryogenic pipeline:
1. Reduce number of the bimetal joints
2.Remove the cooling stations
Bimetal joints
Insulators
To improve the vacuum

## Slide 19: cooling structure of the transition section
Oxygen free copper material is selected for inner cylinder of the liquid helium Dewar.
The press plate and cooling plate will be pre-weld on the inner cylinder to ensure the better cooling effect.

## Slide 20: Summary
In the past ten years,we got very important experience from SSM.
About vacuum and Transition section problem, we are making a new valve box,the design of new one has been completed.
Thanks!