## Slide 1: Hands-on training: Cavity RF measurement
Hongjuan Zheng, Jiyuan Zhai (IHEP, China)
Third Asian School on Superconductivity & Cryogenics for Accelerators (ASSCA2018) hosted by IHEP, China, Dec. 10-16, 2018

## Slide 2: Cavity RF Measurement
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP

## Slide 3: Cavity RF Measurement
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Site (building & lab)
East Hall, Building #6
Prerequisites and reading requirements
Microwave Engineering
Objective
Basic operation of vector network analyzer, field profile measurement of the accelerating mode with perturbation method, Q0 calculation, simulation and measurement, external Q simulation and measurement

## Slide 4: Training Contents
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Vector Network Analyzer usage, calibration and error analysis. (20 min)
Bead pull measurement of the on-axis electric field profiles of TM010 passband modes (50 min)
Question#1: Calculate the max frequency change and phase change of TM010 pi mode due to the Teflon bead perturbation on axis.
Q0 measurement of 2-cell cavity at room temperature (30 min)
Question#2: Calculate Q0 of niobium and copper 650 MHz 2-cell cavity at RT, 77 K, 10 K, 4.5 K, 2 K and 1.8 K.
HOM coupler Qext measurement of TM010 pi, TM110 modes at room temperature (50 min)
Question#3: How to measure Qext of the input coupler and HOM coupler at 2 K? Can we measure Q0 at 2 K only with RF method?

## Slide 5: Bead-pulling Method
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Preparation
Using a non-conducting wire (like a fishing line) to pull a bead (made of dielectric/metallic or ferromagnetic material) through a RF cavity to measure the electromagnetic fields distribution on resonance inside
Principle of bead-pull
Perturbation theory
Form factors of metallic and dielectric ellipsoids
Measurement techniques
Methods and tricks
Bead-pull setup, bead selection, S parameters by VNA

## Slide 6: Principle of Bead-pull
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Cavity performance like resonance frequency f or quality factor Q can be changed by introducing a small perturbation to the resonance EM fields.
Bead-pull measurement involves two types of perturbations:
Small material perturbation, like a small dielectric bead enters a large volume of cavity.
Small cavity volume change, like a small metallic bead enters a large volume of cavity.
Dielectric bead or metallic bead→ perturbation of EM field → frequency and phase changed of EM field→ pulling a metal or dielectric bead in an axial direction → record the frequency shift or phase shift caused by the ball at the corresponding position → get the relative field at that position

## Slide 7: Frequency Shift (1)
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
frequency change due to a metallic bead:
frequency change due to a dielectric bead:
where U is the stored energy in the cavity, a is the radius of the bead, is the relative dielectric constant of the bead material, is the relative magnetic constant of the bead material, is the vacuum dielectric constant, is the vacuum magnetic constant.

## Slide 8: Frequency Shift (2)
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Pulling a bead to measure the cavity Δf(z)/f can measure the sqrt[Ez(r=0,z)]/U:
For example, using a Teflon sphere bead or syringe needle, Hz(r=0,z)≈0 in this case.
Here, Fbead or its alternative factor with multiplication to other constants is called bead form factor which is only defined by the bead geometry and material. So the bead-pull measurement can measure the normalized EM field distribution along the pulling direction which is not cavity frequency and drive power dependent parameter.

## Slide 9: Frequency Shift (3)
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
The perturbation frequency shift Δf in any cavity resonance is found via the change in a S21 measurement according to the relation:
Where QL-1= Q0-1+ Qin-1+ Qout-1 and X= f/f0-f0/f ≈ 2Δf/f0. The Qin and Qout represent the external Q values of input and output probes respectively. From above, we can get the relation between frequency shift with phase shift Φ of S21:

## Slide 10: Beam Pull Measurement System
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
The setup consists of a pulley-mounted driver motor to move the perturbing bead on a dielectric line (‘‘fish line’’) through the cavity.
Coupling to a resonance with a network analyzer can be obtained in cavity across the two beam pipes with antenna.
Antenna used for excitation and pick up
Flange used for bead pulling
Teflon bead
Fish line
Bead pull motor
VNA

## Slide 11: Surface Resistance-NC Cavity
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Conductivity
RF surface resistance
Skin depth
Classic expression:
For normal conductor
Anomalous skin effect:
Dimensionless parameter:
The classical expression for surface resistance applies when αs≦0.016. When αs → ∞
ρl product is a material constant, independent of temperature, as αs ≥ 3, the normal state surface resistance is
R(∞) is the surface resistance in the extreme anomalous limit. The anomalous limit is applicable to a very good conductor, such as copper at microwave frequency and at low temperatures.
For copper: ρl=6.8×10-16Ωm2
For Nb: ρl=6.0×10-16Ωm2
Notes: l is mean free path

## Slide 12: Surface Resistance-SC Cavity
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Surface resistance:
RBCS: BCS resistance
Rres: residual surface resistance
Rres= 5nΩ for 650 MHz 2-cell cavity
f is the rf frequency in GHz
Tc for Nb is 9.2 K
For T<Tc/2,

## Slide 13: Electrical Conductivity for Cu and Nb
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Electrical conductivity for Nb (RRR=300)
Temperature(K) | 2 | 4 | 10 | 80 | 300
σ-Cu (S/m) | 3.16E9 | 3.16E9 | 3.12E9 | 2.55E8 | 5.35E7
σ-Nb (S/m) | / | / | 1.30E9 | 2.00E8 | 7.50E6

## Slide 14: Cavity Quality
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Quality factor
G is independent of the cavity size.
G=284Ω for 650 MHz 2-cell cavity.
R/Q=213Ω for 650 MHz 2-cell cavity.
Geometry constant

## Slide 15: Experimental Set-up of RF Measurement
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
HOM2
HOM1
HOM coupler
Vector Network Analyzer
2-cell cavity
Port1(~input)
Port2(~output)

## Slide 16: RF Measurement of 2-cell Cavity
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Smith chart
Used to decide each port is under coupling, over coupling or matched.
Matched point
VSWR=1
r=1
under coupling
over coupling
Notes: Voltage Standing Wave Ratio

## Slide 17: RF Measurement of 2-cell Cavity
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Different methods used to get the RF parameters of cavity:
Impedance method
Measure resonant VSWR S0, arbitrarily VSWR Sx, resonant frequency f0, arbitrarily frequency fx
Reflection method
Measure f0, 𝑄𝐿, S11(dB), S22(dB)
Transmission method
Measure f0, 𝑄𝐿, S11(dB), S21(dB)
To get the Q0 and Qe for cavity in RT:
First, determine the coupling state of each port.
Second, chose measurement methods.

## Slide 18: RF Measurement-Impedance Method
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
can be achieved by arbitrarily VSWR
Asymmetric VSWR curve
VSWR curve near the resonant frequency

## Slide 19: RF measurement-Reflection Method
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Measure ,,(dB),(dB)
This method has a limitation in that the coupling coefficient is not too small.

## Slide 20: RF measurement-Transmission Method
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Measure ,,(dB),(dB)
This method was used for one port weak coupling.

## Slide 21: References
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
[1] Haipeng Wang, Jiquan Guo, Bead-pulling Measurement Principle and Technique Used for the SRF Cavities at Jlab. US Particle Accelerator School, SRF Practices and Hands-on Measurements, Thomas Jefferson Lab, Newport News, VA 23606 Jan. 19-23, 2015.
[2] Hahn H, Calaga R, Jain P, et al. HOM identification by bead pulling in the Brookhaven ERL cavity. Nuclear Instruments and Methods in Physics Research Section A: Accelerators, Spectrometers, Detectors and Associated Equipment, 2014, 734: 72-78.
[3] Hong-Juan Zheng, Ji-Yuan Zhai, Tong-Xian Zhao, et al. HOMs simulation and measurement results of IHEP02 cavity. Chinese Physics C, 2015, 39(11): 117006.
[4] H. Padamsee, J. Konobloch, T. Hays. RF Superconductivity for Accelerators, Second edition (Strauss GmbH, Mörlenbach, Germany: WILEY-VCH Verlag GmbH & Co. KGaA, 2008), p.44-48, p.77-80, p.85-89.
[5] Material Properties for Engineering Analyses of SRF Cavities, Fermilab Specification: 5500.000-ES-371110.

## Slide 22
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Thank you for your attention!

## Slide 23
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Back slides

## Slide 24: RF Measurement-Impedance Method
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Equivalent circuit with normalized parameter(for Q0 considerations)
cavity
cavity
Equivalent circuit for Qe considerations
Equivalent circuit for QL considerations
These points can be transferred to the smith chart.
Normalized impendence:
unload: half power points:
external : half power points:
loaded : half power points:
For a given measured VSWR at resonance, the VSWR’s that must be achieved at the half-power points for and .

## Slide 25: RF Measurement-Impedance Method
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
Smith chart Q-circle and half power points
Reflection coefficient:
VSWR: S=
Resonant VSWR:()
Arbitrarily VSWR:
VSWR at half-power points:
VSWR at half-power points:

## Slide 26: RF Measurement-Impedance Method
J.Y. Zhai, H.J. Zheng (IHEP, China)
ASSCA2018, IHEP
From above, can be achieved by
Asymmetric VSWR curve
VSWR curve near the resonant frequency