## Slide 1: CEPC所创新基金子课题总结- 数字束流位置测量电子学研制
麻惠洲
加速器中心束测组
2018-12-26

## Slide 2: 主要内容
系统介绍及方案
系统研制进展
系统测试方案
系统初步结论
任务和经费执行情况
项目成员介绍

## Slide 3: 主要内容
系统介绍及方案
系统研制进展
系统测试方案
系统初步结论
任务和经费执行情况
项目成员介绍

## Slide 4: 系统介绍及方案
由于CEPC处于设计阶段，所以数字BPM原型机是基于BII参数来设计并实现的，而数字BPM系统的主要组成部分包括如下几个方面：
BPM探测器部分
HW design
FW design
SW design

## Slide 5: 系统介绍及方案
RTM（AFE）&& AMC（DFE）

## Slide 6: 主要内容
系统介绍及方案
系统研制进展
系统测试方案
系统初步结论
任务和经费执行情况
项目成员介绍

## Slide 7: 数字BPM研制进展主要内容
数字BPM硬件设计
  RTM设计
  AMC设计
数字BPM算法程序设计
数字BPM软件设计

## Slide 8: DBPM.RTM 功能框图
Two LNAs are enough
The place of Attenuator is designed carefully
Attention to the Impedance matching & Noise
-9dB
-3.1dB
Fc=500MHz
BW=10MHz
0dB
to
31dB
NF=0.46dB
Gain=20dB
P1dB=19dBm
IP3=36dBm
NF=2.9dB
Gain=23.75dB
P1dB=18.1dBm
IP3=37dBm
NF=3.39dB
Gain=20dB
P1dB=16.3dBm
IP3=37dBm
Fc=500MHz
BW=10MHz
0dB
to
31dB
LNA
LNA
LNA
Switched Attenuators
-6dB
-3.1dB
Fc=500MHz
BW=10MHz
0dB
to
31dB
NF=0.9dB
Gain=20dB
P1dB=18.1dBm
IP3=37dBm
NF=1.2dB
Gain=20dB
P1dB=16.3dBm
IP3=37dBm
Fc=500MHz
BW=10MHz
LNA
AGC
π
Att
LNA
π
Att
balun

## Slide 9: 数字BPM硬件设计
RTM 电路设计
No cross-calibration
Low Noise design
Low power consumption
Design objective:

## Slide 10: DBPM.RTM S-Parameter Characterization
Receiver S-Parameter Characterization
S21-Parameter Characterization
The performance of band pass filter is good!

## Slide 11: DBPM.RTM Channel to Channel Isolation
-60dBm
-60dBm
--77.473dBm
--78.824dBm
A→B
B→A
6.5dBm→-78.8dBm
6.5dBm→-77.47dBm

## Slide 12: AMC 硬件设计
Power Regulator
Clock Logic
FPGA Logic
ADC Logic
NET Logic
Other Logic…

## Slide 13: Clock Logic
SI571 Configuration（I2C ）
CDCE72010 Configuration （SPI ）

## Slide 14: ADC Logic

## Slide 15: PCB design
Power and GND design

## Slide 16: 数字BPM硬件算法程序设计
基于DDC数字BPM算法的功能框图

## Slide 17: Matlab环境算法模拟

## Slide 18: 算法实现及Chipscope检查

## Slide 19: 其它算法的研究
DFT Algorithm.
Hilbert transformer method.
Time domain dynamic window integrating method.

## Slide 20: 数字BPM软件设计
Software framework include 3 main part:
  Data Acquisition
  Data Analysis
  Data Storage
The Software is built with Python (2.7.10) tools, including module:
  Scipy1.0
  Numpy
  Pandas
  Matplotlib
  Pyepics
  Pyqtgraph

## Slide 21: Data Acquisition Module
  Data Acquisition control logic
  Write/Read CSV file logic
  EPICS IOC Logic
  Data base logic
  Graphical User Interface
  …
SA Data Acquisition function block

## Slide 22: Data Analysis Module
  Data Hilbert Transform and Quadrature Demodulation
  TBT and FA Position calculation
  FFT Analysis.
Note: The function
is designed to verify
the result processed
in FPGA.
Data Analysis function block.

## Slide 23: Data Storage Module
CSV txt format file
RDB Archiver Engineering MySQL database
Two Methods are used in Data Storage:

## Slide 24: 主要内容
系统介绍及方案
系统研制进展
系统测试方案
系统初步结论
任务和经费执行情况
项目成员介绍

## Slide 25: 数字BPM测试方案

## Slide 26: RAW ADC data and Analysis

## Slide 27: TBT（逐圈位置），FA（快速获取位置），SA（闭轨位置） 测试
TBT DATA
FA DATA
SA DATA

## Slide 28: 主要内容
系统介绍及方案
系统研制进展
系统测试方案
系统初步结论
任务和经费执行情况
项目成员介绍

## Slide 29: 数字BPM研制初步结论
经测试，在实验室条件下，在信号发生器产生18dBm点频信号，经1-4功分器，RTM板，在ADC处采样处得到±6000计数时测得TBT，FA，SA的分辨率结果为：
  TBT：达到2µm
  FA： 达到1µm
  SA： 优于0.5µm

## Slide 30: Test with Beam in BEPCII
Have got about 3.17µm resolution for 20 hours with BEPCII real beam.

## Slide 31: 主要内容
系统介绍及方案
系统研制进展
系统测试方案
系统初步结论
任务和经费执行情况
项目成员介绍

## Slide 32: 任务完成情况
BPM前端模拟电子学研究 √
四通道高速ADC精确采样研究 √
BPM数字信号处理算法研究 √
EPICS IOC设计、开发和系统集成 √

## Slide 33: 预算执行情况
经费执行共计：~56.2万
明细 | 经费（万元）
模拟和数字PCB板卡元器件及制版 | 47.4
BPM电子学直流线性电源 | 5.1
FPGA商业开发板(数字BPM算法研究) | 2
BPM分布式定时板卡 | 1.5

## Slide 34: 项目成员及分工情况
项目成员：
  魏书军，叶强，麻惠洲，杜垚垚，刘芳，张醒儿，马宇飞
工作分工：
  总体：魏书军，叶强，麻惠洲
  模拟前端板：杜垚垚，张醒儿
  数字后端板：魏书军，马宇飞
  算法工作：魏书军，刘芳，马宇飞
  软件设计：叶强
  系统测试：叶强，麻惠洲，魏书军，杜垚垚，马宇飞

## Slide 35: 谢谢！