## Slide 1: Status on CBM MRPC @USTC -- R&D on the noise rate
Yongjie Sun State Key Laboratory of Particle Detection and Electronics, China
University of Science and Technology of China

## Slide 2: Outline
Introduction
Simulation and preliminary results
Prototype design
Summary

## Slide 3: CBM TOF
Large area: ~120 m2
Large rate range: 1-25 kHz/cm2
“Streaming” readout: Low noise rate
USTC: MRPC3b (200) + MRPC4 (310)

## Slide 4: CBM/STAR endcap TOF Upgrade
MRPC3b for CBM - STAR: eTOF installation by 2018-12
24 modules ( =72+8 MRPCs) from USTC
- 3 layers
- 8 sectors
- 6912 channels

## Slide 5: Noise rate of batch testing
Noise rate of the strip close to the HV connection much higher than other strips.
Problem

## Slide 6: HV connection
?
What are the main causes of the phenomenon
Guesses
HV connection pad is too close to the strip (2mm)
HV connection pad is too long
Other HV connection pad shapes...

## Slide 7: CST introduction
PCB STUDIO interface
Computer Simulation Technology(CST): the powerful simulation platform for all kinds of electromagnetic field problems and related applications
Website: https://www.cst.com/

## Slide 8: Model_1 simulation
HV pad: 4mm*30mm
Distance between HV pad and strip: 6mm
Inject signal: sine signal
Amplitude: 5mV
Period: 1GHz

## Slide 9: Model_1 simulation results
Amplitude of the closest strip is about 0.17mV,which is about 1/30 of the original signal.
Amplitude of other 2 strips are much smaller than the first strip.
From the side of HV connection pad (the latter result is the same)
P1: first strip P2: second strip P3: third strip

## Slide 10: Model_2 simulation
HV pad: 4mm*15mm
Distance between HV pad and strip:6mm
Input signal: sine signal
Amplitude: 5mV
Period: 1GHz

## Slide 11: Model_2 simulation results
Results:
Amplitude of the closest strip is about 0.1mV,which is about 1/50 of the original signal.
Amplitude of other 2 strips are much smaller than the first strip.
Amplitude of the model_2 is smaller than the model_1.

## Slide 12: Model_3 simulation
No HV pad
Distance between HV pad and strip:6mm
Input signal: sine signal
Amplitude: 5mV
Period: 1GHz

## Slide 13: Model_3 simulation results
Results:
Amplitude of the closest strip is about 0.008mV,which is about 1/625 of the original signal.
Amplitude of other 2 strips are much smaller than the first strip.
Amplitude of the model_3 is the smallest among the 3 models.
HV connection pad has a significant effect on the first strip.

## Slide 14: Model_1modification (position)
2mm
Change HV pad position
As the distance increases, crosstalk decreases significantly!

## Slide 15: Model_1modification (surface resistivity)
Crosstalk and surface resistance are significantly correlated
: surface resistivity
K: graphite conductivity
d: graphite thickness
Parameter setting limitation of the software

## Slide 16: Model_1 modification (pad shape)
Under the same area, the circle looks better than the rectangle.
As the distance increases, crosstalk decreases in these two situation
When the length along the direction of the strip are same, the circle is a bit better than the rectangle.
In fact, The length along the strip is one of the main causes of crosstalk
rectangle pad
circle pad

## Slide 17: Model_1 modification(pad shape)
Area of the pad is not the cause of crosstalk!

## Slide 18: Prototype design -- plan 1
Plan1: Change the position of the rectangular pad(4mm,6mm,8mm)

## Slide 19: Prototype design -- plan 2
Plan2: Keep the area unchanged, the length of the rectangle is reduced by half

## Slide 20: Prototype design -- plan 3
Plan3: use circle pad

## Slide 21: Summary
The noise rate of the strip close to the HV pad was found higher than that of other strips.
We use CST software to study the causes of the phenomenon.
From the results of the simulation, we can get follow conclusion:
As the distance increases, crosstalk decreases significantly.
Crosstalk and surface resistance are significantly correlated.
The length along the strip is one of the main causes of crosstalk.
The circle is a bit better than the rectangle
Based on the simulation results, we have designed some prototypes and the PCB is being produced.
Next step, more simulations will continue and experimental tests will start to verify the simulation results.
Thank you !