## Slide 1: LHAASO-KM2A Muon Detector On-site Test System
MD Group Members
赵世平

## Slide 2: Design of Muon Detectors In KM2A Experiment
(左雄，2015)
The 2.5m thick layer of soil prevents electrons or γ-rays from reaching the detector.
Around 95% purity of muon
will be got.

## Slide 3: Installation of Muon Detectors
1
2
3
4
5
6
7
1~3: build the tank
4: winterize the tank
5: install the PMT
6: inject ultrapure water
7: test the detector

## Slide 4: On-site Test System
基本构造
望远镜
10cm Pb
Scintillator
挑选出Muon事例
WR Switcher
PC
HV
DC
，电子学电源
可控PMT高压源
ED电子学
ED电子学
MD电子学
Readout Crate System

## Slide 5: Off-line Trigger, Events Selection
利用符合得到的事例，采用类似的方法
从MD信号中挑选出缪子事例
事例挑选前
Background
NPE: Number of Photoelectron

## Slide 6: Off-line Trigger
闪烁体
探测器
铅板盒
γ
e
μ
作为能谱输入模拟程序，
结果：
个γ，~700个穿透铅板
个e-，~31个穿透铅板
3533个μ子，全部穿透铅板
缪子的纯度：83%
3m内无e-或者γ的概率90.8%
偶然符合概率不超过9.2%
Shower中
γ，电子和
缪子三种次
级粒子能量
分布
Shower次级
缪子事例3m
范围内的电
子/γ光子数
分布

## Slide 7: 1st Arguments—Number of Photoelectron (NPE)
Trigger
& Events
Selection
Charge Distribution of MD Anode Signals
Fitted by Gauss function
Mean: NPE

## Slide 8: 2nd Argument—Attenuation Time of Signals
探测器内壁反射峰
n — 折射率
— 水中光衰减长度
L — 与探测器几何相关的常数
f — 反射率
平均波形
用指数函数拟合
—
= 174ns

## Slide 9: 50 Muon Detectors
50台探测器的衰减时间和平均光电子数
基本击中在120ns，45NPE附近
信号衰减时间与NPE呈正相关

## Slide 10: Make the Test System More Effective
NPE Distribution (Before Events Selection)
NPE Distribution (After Events Selection)
Mean = 57.05
56.91 = Mean
Using the
telescope
Telescope
（1）too heavy
（2）almost no affection on measuring NPE
（3）no affection on measuring decay time
Exclude the telescope, WR switcher…

## Slide 11: Simplified Test System
PMT
阳极
打拿极
高压
高压模块
PC
Electronics
Network Connection
USB转TTL
单片机
High-Volt
Battery
WiFi
（1）Using batteries, DC 12V
（2）Using wireless network

## Slide 12: Installing the ¼ Array
The new system
will be used to
do the test work