## Slide 1
Huizhou Ma
On behalf of beam diagnostics group of IHEP
Progress on DBPM development in IHEP

## Slide 2: Outline
1
Background
2
The development of Digital BPM electronics
3
The test of DBPM
4
Summary

## Slide 3: Technical features of the new light source
The High Energy Photon Source (HEPS) is designed as an ultra-low emittance ring-based synchrotron radiation light ;
About fourteen beamlines will be constructed in Phase I of the project;
Storage ring circumference: 1360.4m
Energy: 6GeV
Emittance<60pm.rad
Current: >200mA
Construction:2019-2025

## Slide 4: The Digital BPM requirement of HEPS project
HEPS BPM quantities
Storage ring BPM Parameters

## Slide 5: The development of Digital BPM electronics
For the HEPS will have been built several years later, so we develop the DPBM prototype with BEPCII parameters, and DBPM system will be tested in BEPCII.
HW design
  RTM Design
  AMC Design
Test of the DBPM

## Slide 6: RTM Module design
No cross-calibration
Low Noise design
Low power consumption

## Slide 7: DBPM_RTM function block
1、Two LNAs are enough
2、The place of Attenuator is designed carefully
3、Attention to the Impedance matching & Noise
-9dB
-3.1dB
Fc=500MHz
BW=10MHz
0dB
to
31dB
NF=0.46dB
Gain=20dB
P1dB=19dBm
IP3=36dBm
NF=2.9dB
Gain=23.75dB
P1dB=18.1dBm
IP3=37dBm
NF=3.39dB
Gain=20dB
P1dB=16.3dBm
IP3=37dBm
Fc=500MHz
BW=10MHz
0dB
to
31dB
LNA
LNA
LNA
Switched Attenuators
-6dB
-3.1dB
Fc=500MHz
BW=10MHz
0dB
to
31dB
NF=0.9dB
Gain=20dB
P1dB=18.1dBm
IP3=37dBm
NF=1.2dB
Gain=20dB
P1dB=16.3dBm
IP3=37dBm
Fc=500MHz
BW=10MHz
LNA
AGC
π
Att
LNA
π
Att
balun

## Slide 8: S-Parameter Characterization
Receiver S-Parameter Characterization
S21-Parameter Characterization
S21
S11
S12
S22
S21
The performance of band pass filter is good!

## Slide 9: Channel to Channel Isolation
-60dBm
-60dBm
--77.473dBm
--78.824dBm
A→B
B→A
6.5dBm→-78.8dBm
6.5dBm→-77.47dBm

## Slide 10: AMC Function module
Power Logic
Clock Logic
FPGA Logic
ADC Logic
NET Logic
Other Logic…

## Slide 11: RAW ADC data and Analysis
RAW ADC data and analysis

## Slide 12: SA Data testing
RAW ADC data and analysis

## Slide 13: Beam test in BEPCII

## Slide 14: Temperature influence
Thermostatic control Cabinet
Cross-bar
Pilot-Tone
Turn ON Status

## Slide 15: Calibration Bench Test with DBPM
  Input signal: -6dBm
  RAW ADC counts：3k+
  Kx,Ky=5.5
  Step:200µm

## Slide 16: Small-Lot Production of DBPM

## Slide 17: Summary
Digital BPM Electronics hardware has been designed carefully and had been tested in laboratory;
The digital BPM electronics have been tested with the calibration bench and the real beam on BEPCII.
There are still a lot of work to do towards the detail design of BPM.

## Slide 18
Thanks for your attention!