## Slide 1: Beam Instrumentation System of CSNS
Jilei Sun
On behalf of BI group of CSNS, IHEP
MTCA/ATCA Workshop for Research and Industry, Beijing, CHINA
June 23-25, 2019

## Slide 2: Outline
Overview of CSNS project
Introduction of BI system
  Beam Current Monitor, BCM
  Beam Position Monitor, BPM
  Beam Loss Monitor, BLM
  Beam Profile Monitor, Wire Scanner
  Others
Current data acquisition strategy
Summary

## Slide 3: Overview of CSNS project
LINAC service building
RCS service building
Target hall

## Slide 4: Layout of CSNS
50 keV H- Ion source
3 MeV RFQ
80 MeV DTL
1.6 GeV
25 Hz
RCS
Target station and instruments
Peak beam power: 100 kW
50 ~ 500 μs 25 Hz
1 MHz 50% chopping rate
500 ns -> 100 ns

## Slide 5: Beam instrumentation system

## Slide 6: Data Acquisition Strategy
Detectors
Front end electronics
System setup
NI – PXIe cards
AD/DA/DIO
Oscilloscope card
LabVIEW EPICS release
CSS OPI

## Slide 7: Current Monitor
Current Monitor @LINAC
  11 CTs in total
  Magnetic ring from Bergoz
  Magnetic ring from domestic company
Current Monitor @RCS
  3 different types
  DCCT: Commercial product from Bergoz
  SCT/MCT: Self developed sensor
Current Monitor @RTBT
  4 CTs (1 ICT + 3 FCT)
  Magnetic ring from Bergoz

## Slide 8: Beam Current Monitor @LINAC
Beam Current = [Sum(Valid) – Sum(Background)] / n
Valid count
Background count
MACRO pulse measurement, up to 500 μs
Front end electronics + NI PXIe-6358 (16 bit, 1.25 MS/s)

## Slide 9: RCS – DCCT / SCT / MCT
DCCT: Commercial electronics from Bergoz
SCT/MCT: Self developed electronics
200 mA
2 A
20 A
NI PXIe-6358 100 kS/s

## Slide 10: RTBT - CT
Two short pulses, 100 ns each
NI PXIe-5160 oscilloscope card
  1.25 GS/s
  500 MHz band width
  10 bits

## Slide 11: Beam Position Monitor, BPM
LINAC
  29 Strip-line type
  324 MHz RF pulse
  Libera Spark HL, 14 bit
  Bergoz
RCS
  38 Shoe-box
  1 MHz ~ 2 MHz cycling frequency
  Self-developed electronics
  TBT or COD mode
RTBT
  33 Shoe-box
  Libera Spark HL

## Slide 12: Beam Loss Monitor, BLM
156 ion chamber type BLM
  Front end electronics
  NI PXIe-6358 waveform acquisition
  NI PXI-8430 RS232 for HV control
  NI PXI-6713 for threshold configuration for FMPS
15 plastic scintillator + PMT as FBLM
  Libera BLM
  50 MHz bandwidth
  125 MHz / 14 bit

## Slide 13: Beam Profile Monitor, Wire Scanner
20 wire scanners at LINAC&RTBT
7 multi-wire at the injection & dump
  Stepper motor driven
  NI PXI-7330/7358 motor control
  NI PXIe-6358 do the DAQ after front end
Wire Scanner
Multi-Wire
Front end Electronics

## Slide 14: Profile Monitor @RCS - IPM
Ionization Profile Monitor
  64 channel signals
  ≥ 14 bit
  20~100 MHz sampling rate
  Memory size: ≥ 100 MB / channel

## Slide 15: Summary of the DAQ requirement
Detector | Type of Signal | No. of channels | Sampling Rate | Resolution
CT | 50~500 μs, 25 Hz 500~100 ns, 25Hz | 18 | 1.23 MHz@LINAC 100 kHz@RCS 1.25 GHz@RTBT | 16 bit
BPM | 324 MHz 1~2 MHz 100 ns, 25Hz | 100*4 | 125 MHz@LINAC 60 MHz@RCS 125 MHz@RTBT | ≥14 bit
BLM | 50~500 μs, 25 Hz 500~100 ns, 25Hz | 156+15 | 500 kHz@LINAC 50 kHz@RCS 100 kHz@RTBT | 16 bit
Profile | 100 μs, 1Hz 500~100 ns, 25 Hz | ~400 | 1 MHz@LINAC ≥20 MHz@RCS 1 MHz@RTBT | 16 bit
FCT | 324 MHz | 13 | 100 MHz | 14 bit

## Slide 16: Shortage of current system
17 PXI Windows host computers
6 times OS failures in last one year
NI real time OS: Phar Lap ETS
  I/O control
  Data storage
  EPICS configuration
Chart: 
| LRBTCT | R2BLM | R4BLM | RTBLM | TargetMWS
 | 2 | 1 | 1 | 1 | 1

## Slide 17: Summary
CSNS running @50 kW stably, opening to users already. The goal is have ~80 kW, by the end of this year.
DAQ system of our beam instrumentation nearly all based on NI PXI/PXIe platform. Performance acceptable, and we are keep improving and updating on software.
PXI based EST real time OS is under developing.

## Slide 18
Thanks for your attention！