## Slide 1: Performance of the GridPix quad TPC readout
Cornelis Ligtenberg, Y. Bilevych, K. Desch, H. van der Graaf, M. Gruber, F. Hartjes, K. Heijhoff,
J. Kaminski, P.M. Kluit, N. van der Kolk, G. Raven, T. Schiffer, J. Timmermans
The 2019 International Workshop on the High Energy Circular Electron Positron Collider, November 18, Beijing

## Slide 2: Introduction and outline
The goal is to develop a pixel readout for a large TPC
A large TPC with a 55 µm × 55 µm pixel readout
First Timepix3 based GridPix test beam (2017)
Quad module performance from test beam (2018)
Development of 8 quad detector (2019-2020)
Picture ILC TDR

## Slide 3: GridPix technology
GridPix is a type of micro-pattern gaseous TPC readout
The GridPix based on a Timepix3
  55 µm × 55 µm pixels
  Digital simultaneous registration of Time of Arrival (1.56 ns) and Time over Threshold
  An aligned Aluminium amplification grid is added by photolithographic postprocessing techniques
Single ionization electrons are detected with high efficiency
  The maximum possible information from a track is acquired
  dE/dx by cluster counting
50 µm
35 µm
55 µm
1 µm

## Slide 4: Simulation of ILD TPC with pixel readout
To study the performance of a large pixelised TPC, the pixel readout was implemented in the full ILD DD4HEP (Geant4) simulation
Changed the existing TPC pad readout to a pixel readout
Adapted Kalman filter track reconstruction to pixels
Pads
Pixels
22 electrons / hit
200 hits / track
1 electron / hit
10 000 hits / track
Picture by Cornell
6 mm
55 µm
50 GeV muon track with pixel readout
Continuous
3D tracking

## Slide 5: Performance of a GridPix TPC at ILC
Momentum resolution of the TPC for a 50 GeV muon
From full simulation, momentum resolution can be determined
Momentum resolution is 15% better (with realistic coverage and delta’s)
Realistic tiling with quad module
58% coverage

## Slide 6: Single chip results (2017)
A GridPix based on the Timepix3 chip was reliably operated in a test beam setup with 2.5 GeV electrons at ELSA (Bonn)
T2K gas and Edrift = 280 V/cm, Vgrid = -350 V
The resolution is primarily limited by diffusion
Systematic uncertainties are low: < 10 µm in plane
Energy loss resolution (dE/dx) by electron counting is 4.1 % per meter
Published paper on this testbeam doi:10.1016/j.nima.2018.08.012
Electron
MIP
Hit resolution
B T

## Slide 7: The quad module
A four chip module sized 39.6 mm 28.38 mm
The quad module has all services under the active area
  Can be tiled to cover arbitrarily large areas.
Area for connections IO was minimized
  Maximises active area ( 68.9% )
To maintain a homogenous electric field wire bonds are covered by a central guard
High precision < 20 μm mounting of the chips and guard
Cooling channels

## Slide 8: Test beam measurements (2018)
2.5 GeV electrons at the ELSA accelerator in Bonn, Germany
T2K gas with Edrift = 280 V/cm, Vgrid = -350 V
Events are triggered by a scintillating plane
6 plane mimosa telescope with 18.4 μm × 18.4 μm sized pixels
180 o

## Slide 9: Time walk correction
Time walk error: time of arrival depends on signal amplitude
Time walk can be corrected using Time over Threshold (ToT) as a measure for signal strength
First order correction fitted and applied:
Distribution of residuals becomese more Gaussian after the time walk correction

## Slide 10: Hit resolution in the drift direction
Single hit resolution in drift direction
Depends on
from fit
Diffusion from fit
Because of a large time walk error in hits with a low signal strength, an additional ToT cut ( > 0.60 µs) was imposed

## Slide 11: Hit resolution in the pixel (precision) plane
Single hit resolution in pixel (precision) plane:
Depends on:
pixel size 55 µm
Diffusion from fit
B T
Scattering on guard
Note that:
A hit resolution of 250 µm is 25 µm for a 100-hit track ( 1 cm track length)
At 4 T , expected
At 2 T , expected

## Slide 12: Deformations in the pixel (precision) plane
Investigation of systematic deviations over the pixel plane
Each bin displays mean of residuals from 4 × 4 pixels
Primarily due to electric field distortions
Correction of deformations with 4 fitted Cauchy functions per chip:
Without correction

## Slide 13: Deformations in the pixel (precision) plane
Investigation of systematic deviations over the pixel plane
Primarily due to electric field distortions
Each bin displays mean of residuals from 4 × 4 pixels
Correction of deformations with 4 fitted Cauchy functions per chip:
With correction

## Slide 14: Deformations in pixel plane and drift direction
Each bin displays mean of residuals from 4 × 4 pixels
The RMS in the center of the chip is 9 µm (pixel plane after correction) and 14 µm (drift direction), which indicates small systematic errors
Corrected x deformations
Pixel (precision) plane
z deformations
Drift direction

## Slide 15: 8 quad module development
8 quad test box with (32 chips)
Simultaneous read out through one SPIDR board using a data concentrator
Field wires added to improve electric field, and reduce deformations
New field wires

## Slide 16: 8 quad module development
Laser test indicate a reduction in electric field deformations with field wires
Early 2020 test beam planned at DESY with 1 T magnetic field
Preliminary
Uncorrected residuals from quad test beam
Uncorrected residuals from laser test with field wires

## Slide 17: A GridPix TPC at the CEPC?
A GridPix TPC can deal with the high beam rates at the CEPC
The CEPC with L = 32 1034 cm-2 s-1 will produce Z bosons at 10 kHz
Link speed of Timepix3 (in Quad) is 80 Mbps: 2.6 MHits/s per 1.41 × 1.41 cm2
Excellent time resolution: time stamping of tracks < 1.2 ns
Power consumption 2W/chip depending on hit rate
  No power pulsing possible at the CEPC
  Good cooling is important
Picture IHEP
Rough estimation
Rough estimation

## Slide 18: A GridPix TPC at the CEPC?
Ions are not limiting performance in the 240 GeV (Higgs) runs
However, the number of Ions in the high luminosity 91 GeV run might be high
Rough estimations at L = 32⋅1034 cm-2 s-1 indicate primary ionisation causes accumulated charge at an ILC250 level distortions < 5 µm (see backup slide)
Ion backflow (IBF) can give a lot of additional charge, so IBF must be controlled
Measuring IBF for Gridpix is a priority, expected
Gating can greatly reduce IBF
  At CEPC gating is possible because:
  max drift time of 30 µs < average Z interval 100 µs (10 kHz)
  Will cause some leveling due to dead time
Rough estimation

## Slide 19: Conclusions
A quad module with four Timepix3 based GridPixes has been designed and built
  The resolution is limited by diffusion
  Systematic uncertainties are small: 9 µm (pixel plane) and 14 µm (drift direction)
A 8 quad detector with 32 chips is realized and is under investigation
Simulations show an improvement in momentum resolution of a pixel TPC readout over a pad readout
A GridPix pixel TPC is an interesting option for an experiment at the CEPC:
  High precision tracking in the transverse and longitudinal planes
  dE/dx by electron and cluster counting
  Excellent two track resolution
  Digital readout that can deal with high rates

## Slide 20: Backup

## Slide 21: Ions in CEPC TPC
Rough estimations at L = 35⋅1035 cm-2 s-1 indicate primary ionisation at a ILC250 level < 5 µm distortions (This equals 8 µm with IBF = 1?) See Arai Daisuke
Simulation from CEPC CDR with Gain × IBF = 5 and L = 17⋅1034 cm-2 s-1 < 40 µm distortions ( This equals 16 µm at Gain × IBF = 1 and L = 32⋅1034 cm-2 s-1 )
FCCee/TLEP studies at Gain × IBF = 1 and 16.8 kHz hadronic Zs by Philippe Schwemling < 22 µm distortions
Rough esitimation of primary ionisation
10 kHz Z event rate
500 ms will accumulate 5000 Z events
20 tracks / Z event and 10 000 e / track will make 108 ions in volume
Volume is ~4 107 resulting in 25 e/cm3
Similar to ILC250 accumulated charge
Rough estimation

## Slide 22: Ion backflow
Ion backflow also needs to be controlled as this might lead to large distortions
Old measurements from GridPix (thesis M. Chefdeville) indicate backflow can be reduced to per mil level
New measurements are a priority

## Slide 23: Timepix3 pixel chip
256 × 256 pixels with 55 µm × 55 µm pitch
Sensitive area of 14.1 mm × 14.1 mm
TDC with 640 MHz clock, resulting in a 1.56 ns time resolution
Per pixel simultaneous measurement of arrival time (ToA) and signal amplitude (ToT)
Readout using SPIDR
Power consumption of 2W depending on hit rate
  good cooling is important
Wafer post-processed at IZM Berlin
Sensitive area
2+3 mm
14.1 mm
14.1 mm
64 bit data packets

## Slide 24: Run parameters and selections
Used T2K (Ar:CF4:iC4H10 95:3:2) gas with a water vapor contamination
  Drift speed 54.6 µm/ns (59.0 µm/ns expected by magboltz)
Most probable number of hits per 27.5 mm was 146 (225 expected)
  This is due to the low effective grid voltage and possibly read out problems
Use a stringent selection to get clean tracks

## Slide 25: Resolution of quad module
Determine overall accuracy of a track position measurement by comparing the quad track with the telescope track
Subtract a background of unrelated tracks
Error contributions:
Statistical error using hit resolution
Systematic errors from RMS in pixel plane and drift direction
Multiple scattering contribution from simple Monte Carlo simulation
In the end, an unidentified contribution remains
Observed standard deviation | 41 µm
Statistical errors | 25 µm
Systematic errors in the pixel plane and drift direction | 19 µm
Multiple scattering | 22 µm
Unidentified systematic error | 14 µm

## Slide 26: Production of GridPixes
Cleaning
Deposition of Protection layer
SU-8 covering
Exposure with mask
Aluminium layer is deposited
Another layer of photoresist is applied, exposer with a mask creates a hole pattern, and the holes are chemically etched
The wafer is diced
The unexposed SU-8 is resolved
Thesis Stergios Tsigaridas, Next Generation GridPix

## Slide 27: Motivation for a pixelised TPC
Improved dE/dx by cluster counting
Improved measurement of low angle tracks
Improved double track seperation
Much reduced hodoscope effect
Lower occupancy in high rate environments
Fully digital read out