## Slide 1: KM2A PMT testing platform
FENG Cunfeng,
LI Chaoju, SUN Yansheng
Shandong University
THE 2nd WORKSHOP OF AIR SHOWER DETECTION AT HIGH ALTITUDES
IHEP, Beijing, 17~19, Feb. 2011

## Slide 2: Outline
LHAASO KM2A PMT
PMT testing scheme and instrument
Primary testing results
Summary

## Slide 3: LHAASO KM2A PMT
KM2A detector unit:
  Each detector unit consist of 16 plastic tiles and one PMT.
  PMT couple to plastic tiles by around 100 optical fibers.
  Dynamical range: 1~3000 charge particles
  Total 5,000 units, means 5,000 PMTs
Characteristics of PMT for KM2A
  anode linearity range: ~100mA within 5%
  spatial uniformity: <10% in the center region of the window
  TTS(transit time spread): ~1ns
  Low noise
  Gain: 106

## Slide 4: Schematic diagram of the testing system
PC
WinXP+Labview
VME
HV
SY1527
dark box
QDC
LTD
scale
Pulse generator
Diffusion filter
LED1
ps pulse laser
amp
Aging box1
Aging box2
Aging box3
LED
Digital motor
fan in/out
15 PMTs
15 PMTs
LED2
Oscilloscope
The whole system is controlled by the central compute

## Slide 5: Instrument
VME DAQ system
HV power supply
  SY1527 +70 channel output
Scanning & testing dark box
  scan 16 PMTs simultaneously
Pulse generator
  8 output channels
Oscilloscope
  band width: 500MHz
Picosecond pulse laser, LED driver

## Slide 6: Control panel
Developed with LabView
Main panel
Function panel
  DAQ(VME control)
  Pulse generator
  HV power supply
  Scan motor
  Oscilloscope

## Slide 7: Linearity measurement
Two LEDs used: LED1, LED2
  LED driver amplitude: 0~20V
  LED driver width: 20ns
  LED driver triggered by pulse generator
LED1 and 2 light alternately and together
linearity=[Iled1+led2-(Iled1+Iled2)]/Iled1+led2*100%

## Slide 8: Uniformity measurement
Optical fiber scan the PMT window in 2D.
  Step pace:1mm
  Control by digital motor
Central computer communicate with digital motor, LED driver and DAQ system

## Slide 9: Rise time, TTS
Light source : picosecond pulsed light (laser diode)
  trigger by the generator
  Width: ~100ps
Oscilloscope used to measure the rise time, TTS
  The measurement results are read to central PC with ethernet .
  One channel input the trigger .
  3 channels for PMTs signal.
rise time read from oscilloscope directly
TTS(transit time spread):
  Read the PMT signal and laser trigger signal wave shape
  Transit time=Tsignal-Ttrigger-Tdelay

## Slide 10: Gain & Single photoelectron peak
Gain
  XP2622
  HV: 1400-2200V
Single photoelectron peak
  XP2622
  HV: 1800V

## Slide 11: Summary
The PMT testing platform almost ready
The platform is easy to use
  One push, all results will automatic present
    Except the oscilloscope, need switch manually.
The platform is speedy
  One batch 15 PMTs, testing simultaneously
  One or two batch per day.
  One year or half year will finish all 5,000 tubes.

## Slide 12
Thanks for your attention!